## Instruction
Instructions for analyzing GeneRead Library Quant Kit or Array (Illumina) results with this file:
Generally, only change data in yellow or orange cells. Gray and white cells contain formulas for calculation or results. Please do not change them.
1. Raw data worksheet: Copy and Paste the CT values from the real-time PCR instrument runs of your experimental samples to the yellow section of the "Raw data" worksheet starting at cell B2.
2. Sample library dilution: Input your sample library dilution fold (in orange) if you use different dilution fold other than recommended ones in handbook.
3. Library size: Input your library size (in orange) if your library size is significantly different from 426 bp (+/-25bp).This will be used for size adjustment to account for the difference in size between size of sample library and the DNA standard (426 bp amplicon).
4. QC report: All NTC should be more than 29. Slope should be within –3.1 to –3.6. Standard deviation for triplicates should be <0.4.
5. Results: Final library concentrations (in pM) are displayed, from which you can calculate the dilution fold to dilute your original libraries for cluster generation.
6. Calculations: To view the formulas and equations used to calculate the QC data and results, you may review but not alter the "Calculations" worksheet.
.

## Raw data
Well | raw CT
1 | 11.4267
2 | 11.4267
3 | 11.4267
4 | 14.7
5 | 14.7
6 | 14.7
7 | 18
8 | 18
9 | 18
10 | 21.3
11 | 21.3
12 | 21.3
13 | 24.7
14 | 24.7
15 | 24.7
16 | 40
17 | 40
18 | 40
19 | 11.4267
20 | 11.4267
21 | 11.4267
22 | 14.7
23 | 14.7
24 | 14.7
25 | 11.4267
26 | 11.4267
27 | 11.4267
28 | 14.7
29 | 14.7
30 | 14.7
31 | 11.4267
32 | 11.4267
33 | 11.4267
34 | 14.7
35 | 14.7
36 | 14.7
37 | 11.4267
38 | 11.4267
39 | 11.4267
40 | 14.7
41 | 14.7
42 | 14.7
43 | 11.4267
44 | 11.4267
45 | 11.4267
46 | 14.7
47 | 14.7
48 | 14.7
49 | 11.4267
50 | 11.4267
51 | 11.4267
52 | 14.7
53 | 14.7
54 | 14.7
55 | 11.4267
56 | 11.4267
57 | 11.4267
58 | 14.7
59 | 14.7
60 | 14.7
61 | 11.4267
62 | 11.4267
63 | 11.4267
64 | 14.7
65 | 14.7
66 | 14.7
67 | 11.4267
68 | 11.4267
69 | 11.4267
70 | 14.7
71 | 14.7
72 | 14.7
73 | 11.4267
74 | 11.4267
75 | 11.4267
76 | 14.7
77 | 14.7
78 | 14.7
79 | 11.4267
80 | 11.4267
81 | 11.4267
82 | 14.7
83 | 14.7
84 | 14.7
85 | 11.4267
86 | 11.4267
87 | 11.4267
88 | 14.7
89 | 14.7
90 | 14.7
91 | 11.4267
92 | 11.4267
93 | 11.4267
94 | 14.7
95 | 14.7
96 | 14.7
97
98
99
100

## Sample library dilution
Well1 | Well2 | Well3 | Sample name | Dilution fold | X | pM
1 | 2 | 3 | std1 |  | 1.301 | 20
4 | 5 | 6 | std2 |  | 0.30103 | 2
7 | 8 | 9 | std3 |  | -0.69897 | 0.2
10 | 11 | 12 | std4 |  | -1.699 | 0.02
13 | 14 | 15 | std5 |  | -2.699 | 0.002
16 | 17 | 18 | NTC
19 | 20 | 21 | sample library 1 dilution 1 | 2000
22 | 23 | 24 | sample library 1 dilution 2 | 20000
25 | 26 | 27 | sample library 2 dilution 1 | 2000
28 | 29 | 30 | sample library 2 dilution 2 | 20000
31 | 32 | 33 | sample library 3 dilution 1 | 2000
34 | 35 | 36 | sample library 3 dilution 2 | 20000
37 | 38 | 39 | sample library 4 dilution 1 | 2000
40 | 41 | 42 | sample library 4 dilution 2 | 20000
43 | 44 | 45 | sample library 5 dilution 1 | 2000
46 | 47 | 48 | sample library 5 dilution 2 | 20000
49 | 50 | 51 | sample library 6 dilution 1 | 2000
52 | 53 | 54 | sample library 6 dilution 2 | 20000
55 | 56 | 57 | sample library 7 dilution 1 | 2000
58 | 59 | 60 | sample library 7 dilution 2 | 20000
61 | 62 | 63 | sample library 8 dilution 1 | 2000
64 | 65 | 66 | sample library 8 dilution 2 | 20000
67 | 68 | 69 | sample library 9 dilution 1 | 2000
70 | 71 | 72 | sample library 9 dilution 2 | 20000
73 | 74 | 75 | sample library 10 dilution 1 | 2000
76 | 77 | 78 | sample library 10 dilution 2 | 20000
79 | 80 | 81 | sample library 11 dilution 1 | 2000
82 | 83 | 84 | sample library 11 dilution 2 | 20000
85 | 86 | 87 | sample library 12 dilution 1 | 2000
88 | 89 | 90 | sample library 12 dilution 2 | 20000
91 | 92 | 93 | sample library 13 dilution 1 | 2000
94 | 95 | 96 | sample library 13 dilution 2 | 20000

## Library size
Sample Name | Library size
sample library 1 | 280
sample library 2 | 280
sample library 3 | 280
sample library 4 | 280
sample library 5 | 280
sample library 6 | 280
sample library 7 | 280
sample library 8 | 280
sample library 9 | 280
sample library 10 | 280
sample library 11 | 280
sample library 12 | 280
sample library 13 | 280

## QC report
[6 empty] | Well1 | Well2 | Well3 | Sample name | Dilution fold | Value1 | Value2 | Value3 | Ave CT (Y) | STDEV
NTC QC | [5 empty] | 1 | 2 | 3 | std1 |  | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15
[6 empty] | 4 | 5 | 6 | std2 |  | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15
Well1 | Well2 | Well3 | Name | Ave CT(Y) | Pass/Fail | 7 | 8 | 9 | std3 |  | 18 | 18 | 18 | 18 | 0
16 | 17 | 18 | NTC | 40 | Pass | 10 | 11 | 12 | std4 |  | 21.3 | 21.3 | 21.3 | 21.3 | 0
[6 empty] | 13 | 14 | 15 | std5 |  | 24.7 | 24.7 | 24.7 | 24.7 | 0
[6 empty] | 16 | 17 | 18 | NTC |  | 40 | 40 | 40 | 40 | 0
[6 empty] | 19 | 20 | 21 | sample library 1 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15
Slope QC | [5 empty] | 22 | 23 | 24 | sample library 1 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15
[6 empty] | 25 | 26 | 27 | sample library 2 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15
| slope | intercept | RSQ | Efficiency | Pass/Fail | 28 | 29 | 30 | sample library 2 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15
| -3.3147 | 15.7085 | 0.99995 | 1.003 | Pass | 31 | 32 | 33 | sample library 3 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15
[6 empty] | 34 | 35 | 36 | sample library 3 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15
[6 empty] | 37 | 38 | 39 | sample library 4 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15
[6 empty] | 40 | 41 | 42 | sample library 4 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15
[6 empty] | 43 | 44 | 45 | sample library 5 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15
[6 empty] | 46 | 47 | 48 | sample library 5 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15
[6 empty] | 49 | 50 | 51 | sample library 6 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15
[6 empty] | 52 | 53 | 54 | sample library 6 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15
[6 empty] | 55 | 56 | 57 | sample library 7 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15
[6 empty] | 58 | 59 | 60 | sample library 7 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15
[6 empty] | 61 | 62 | 63 | sample library 8 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15
[6 empty] | 64 | 65 | 66 | sample library 8 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15
[6 empty] | 67 | 68 | 69 | sample library 9 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15
[6 empty] | 70 | 71 | 72 | sample library 9 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15
[6 empty] | 73 | 74 | 75 | sample library 10 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15
[6 empty] | 76 | 77 | 78 | sample library 10 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15
[6 empty] | 79 | 80 | 81 | sample library 11 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15
[6 empty] | 82 | 83 | 84 | sample library 11 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15
[6 empty] | 85 | 86 | 87 | sample library 12 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15
[6 empty] | 88 | 89 | 90 | sample library 12 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15
[6 empty] | 91 | 92 | 93 | sample library 13 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15
[6 empty] | 94 | 95 | 96 | sample library 13 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15

## Results
Well1 | Well2 | Well3 | Sample name | pM | Ave CT (Y)
1 | 2 | 3 | std1 | 20 | 11.4267
4 | 5 | 6 | std2 | 2 | 14.7
7 | 8 | 9 | std3 | 0.2 | 18
10 | 11 | 12 | std4 | 0.02 | 21.3
13 | 14 | 15 | std5 | 0.002 | 24.7
16 | 17 | 18 |  |  | 40
19 | 20 | 21 | sample library 1 dilution 1 | 59574.4098 | 11.4267
22 | 23 | 24 | sample library 1 dilution 2 | 61309.755 | 14.7
25 | 26 | 27 | sample library 2 dilution 1 | 59574.4098 | 11.4267
28 | 29 | 30 | sample library 2 dilution 2 | 61309.755 | 14.7
31 | 32 | 33 | sample library 3 dilution 1 | 59574.4098 | 11.4267
34 | 35 | 36 | sample library 3 dilution 2 | 61309.755 | 14.7
37 | 38 | 39 | sample library 4 dilution 1 | 59574.4098 | 11.4267
40 | 41 | 42 | sample library 4 dilution 2 | 61309.755 | 14.7
43 | 44 | 45 | sample library 5 dilution 1 | 59574.4098 | 11.4267
46 | 47 | 48 | sample library 5 dilution 2 | 61309.755 | 14.7
49 | 50 | 51 | sample library 6 dilution 1 | 59574.4098 | 11.4267
52 | 53 | 54 | sample library 6 dilution 2 | 61309.755 | 14.7
55 | 56 | 57 | sample library 7 dilution 1 | 59574.4098 | 11.4267
58 | 59 | 60 | sample library 7 dilution 2 | 61309.755 | 14.7
61 | 62 | 63 | sample library 8 dilution 1 | 59574.4098 | 11.4267
64 | 65 | 66 | sample library 8 dilution 2 | 61309.755 | 14.7
67 | 68 | 69 | sample library 9 dilution 1 | 59574.4098 | 11.4267
70 | 71 | 72 | sample library 9 dilution 2 | 61309.755 | 14.7
73 | 74 | 75 | sample library 10 dilution 1 | 59574.4098 | 11.4267
76 | 77 | 78 | sample library 10 dilution 2 | 61309.755 | 14.7
79 | 80 | 81 | sample library 11 dilution 1 | 59574.4098 | 11.4267
82 | 83 | 84 | sample library 11 dilution 2 | 61309.755 | 14.7
85 | 86 | 87 | sample library 12 dilution 1 | 59574.4098 | 11.4267
88 | 89 | 90 | sample library 12 dilution 2 | 61309.755 | 14.7
91 | 92 | 93 | sample library 13 dilution 1 | 59574.4098 | 11.4267
94 | 95 | 96 | sample library 13 dilution 2 | 61309.755 | 14.7 | Notes:
If both dilutions result in CT values within the dynamic range of standards, use pM values from the most concentrated library dilution. If one dilution only falls within the dynamic range of standards, choose the pM value from the dilution that yields a CT value within the dynamic range of standards.
If CT values from both dilutions fall out of range (low CT), the library is too concentrated. We recommend checking the concentration using BioAnalyzer or QIAxcel and diluting the library to about 1 nM, and repeating library quantification using the GeneRead Library Quant System.
If CT values from both dilutions fall out of range (high CT), the library concentration is too low and not good for proceeding to sequencing. Library construction has to be repeated.

## Calculation
Well | raw CT | Well1 | Well2 | Well3 | Sample name | Dilution fold | Value1 | Value2 | Value3 | Ave CT (Y) | STDEV | X | pM
1 | 11.4267 | 1 | 2 | 3 | std1 |  | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15 | 1.301 | 20
2 | 11.4267 | 4 | 5 | 6 | std2 |  | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15 | 0.30103 | 2
3 | 11.4267 | 7 | 8 | 9 | std3 |  | 18 | 18 | 18 | 18 | 0 | -0.69897 | 0.2
4 | 14.7 | 10 | 11 | 12 | std4 |  | 21.3 | 21.3 | 21.3 | 21.3 | 0 | -1.699 | 0.02
5 | 14.7 | 13 | 14 | 15 | std5 |  | 24.7 | 24.7 | 24.7 | 24.7 | 0 | -2.699 | 0.002
6 | 14.7 | 16 | 17 | 18 | NTC |  | 40 | 40 | 40 | 40 | 0
7 | 18 | 19 | 20 | 21 | sample library 1 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15 | 1.2918 | 59574.4098
8 | 18 | 22 | 23 | 24 | sample library 1 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15 | 0.30425 | 61309.755
9 | 18 | 25 | 26 | 27 | sample library 2 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15 | 1.2918 | 59574.4098
10 | 21.3 | 28 | 29 | 30 | sample library 2 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15 | 0.30425 | 61309.755
11 | 21.3 | 31 | 32 | 33 | sample library 3 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15 | 1.2918 | 59574.4098
12 | 21.3 | 34 | 35 | 36 | sample library 3 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15 | 0.30425 | 61309.755
13 | 24.7 | 37 | 38 | 39 | sample library 4 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15 | 1.2918 | 59574.4098
14 | 24.7 | 40 | 41 | 42 | sample library 4 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15 | 0.30425 | 61309.755
15 | 24.7 | 43 | 44 | 45 | sample library 5 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15 | 1.2918 | 59574.4098
16 | 40 | 46 | 47 | 48 | sample library 5 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15 | 0.30425 | 61309.755
17 | 40 | 49 | 50 | 51 | sample library 6 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15 | 1.2918 | 59574.4098
18 | 40 | 52 | 53 | 54 | sample library 6 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15 | 0.30425 | 61309.755
19 | 11.4267 | 55 | 56 | 57 | sample library 7 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15 | 1.2918 | 59574.4098
20 | 11.4267 | 58 | 59 | 60 | sample library 7 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15 | 0.30425 | 61309.755
21 | 11.4267 | 61 | 62 | 63 | sample library 8 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15 | 1.2918 | 59574.4098
22 | 14.7 | 64 | 65 | 66 | sample library 8 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15 | 0.30425 | 61309.755
23 | 14.7 | 67 | 68 | 69 | sample library 9 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15 | 1.2918 | 59574.4098
24 | 14.7 | 70 | 71 | 72 | sample library 9 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15 | 0.30425 | 61309.755
25 | 11.4267 | 73 | 74 | 75 | sample library 10 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15 | 1.2918 | 59574.4098
26 | 11.4267 | 76 | 77 | 78 | sample library 10 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15 | 0.30425 | 61309.755
27 | 11.4267 | 79 | 80 | 81 | sample library 11 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15 | 1.2918 | 59574.4098
28 | 14.7 | 82 | 83 | 84 | sample library 11 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15 | 0.30425 | 61309.755
29 | 14.7 | 85 | 86 | 87 | sample library 12 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15 | 1.2918 | 59574.4098
30 | 14.7 | 88 | 89 | 90 | sample library 12 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15 | 0.30425 | 61309.755
31 | 11.4267 | 91 | 92 | 93 | sample library 13 dilution 1 | 2000 | 11.4267 | 11.4267 | 11.4267 | 11.4267 | 2.1756e-15 | 1.2918 | 59574.4098
32 | 11.4267 | 94 | 95 | 96 | sample library 13 dilution 2 | 20000 | 14.7 | 14.7 | 14.7 | 14.7 | 2.1756e-15 | 0.30425 | 61309.755
33 | 11.4267
34 | 14.7 | slope | intercept | RSQ | efficiency
35 | 14.7 | -3.3147 | 15.7085 | 0.99995 | 1.003
36 | 14.7
37 | 11.4267
38 | 11.4267
39 | 11.4267
40 | 14.7
41 | 14.7
42 | 14.7
43 | 11.4267
44 | 11.4267
45 | 11.4267
46 | 14.7
47 | 14.7
48 | 14.7
49 | 11.4267
50 | 11.4267
51 | 11.4267
52 | 14.7
53 | 14.7
54 | 14.7
55 | 11.4267
56 | 11.4267
57 | 11.4267
58 | 14.7
59 | 14.7
60 | 14.7
61 | 11.4267
62 | 11.4267
63 | 11.4267
64 | 14.7
65 | 14.7
66 | 14.7
67 | 11.4267
68 | 11.4267
69 | 11.4267
70 | 14.7
71 | 14.7
72 | 14.7
73 | 11.4267
74 | 11.4267
75 | 11.4267
76 | 14.7
77 | 14.7
78 | 14.7
79 | 11.4267
80 | 11.4267
81 | 11.4267
82 | 14.7
83 | 14.7
84 | 14.7
85 | 11.4267
86 | 11.4267
87 | 11.4267
88 | 14.7
89 | 14.7
90 | 14.7
91 | 11.4267
92 | 11.4267
93 | 11.4267
94 | 14.7
95 | 14.7
96 | 14.7