## Instructions
Using the QIAseq® Targeted DNA Panels with UDIs Template for the Local Run Manager (LRM) v3 Excel® File from Illumina®
Release date: December 2021 ¦ Version: 1.00 ¦ Based on the QIAseq Targeted DNA Panel Handbook, HB-2168-004
Descriptions of Tab Functions Note: Generally, only change data in highlighted yellow cells. 1. Read Allocation tab provides guidance on how much coverage is needed and how many reads to allocate to each library and many samples can be sequenced together depending on the panel size, input amount, and Variant Allele Frequency (VAF). Fill in the Primer number in cell Q18 and DNA coverage depending on input and VAF in cell S33. 2. Reagent Tracking tab is where you would fill in the Lot number, Date opened, and Expiry date of each reagent used in the experiment. 3. Protocol tab provides guidance on targeted DNA enrichment for Illumina instruments with Local Run Manager. Fill in the Sample ID, Panel library, DNA conc. (ng/µl), and the desired ng per sample input amount. The dilutions will then be automatically calculated for a total of DNA (ng) that is specified in well F33 to go directly into the Fragmentation, End-Repair, and A-Tailing Reaction. From the drop-down menus under Well ID, starting at H35, select which Well ID will be used for each of the samples. The Index information for both the P50n and P70n will automatically autopopulate. Do not duplicate indices. Cells in Column H will display red font if indices are duplicated, for a warning. Also, note that the master mixes are all calculated with at least a 10% excess. 4. Index Plates tab describes and illustrates the positions of the indices within the index plates. 5. Dilute Lib tab. Copy-paste the concentrations in pM from the Results tab, starting at cell B2, in either the QIAseq Library Quant Assay or Array Kit_96_DataAnalysis_2016.xlsx or QIAseq Library Quant Assay or Array Kit_384_DataAnalysis_2016_HT 96 samples.xlsx to the yellow section of the Dilute Lib tab, starting at cell C14. After this, dilution for the samples to 4 nM will be automatically calculated. Add 3 µl of the library to the appropriate amount of water or HT1 Buffer. If your samples are too dilute, you may change cell E13 to be less than 4 nM, but you will need to do the remainder of the calculations. 6.–8. Dilut and Denat Lib MiSeq tab describes how to dilute the NaOH appropriately, how to denature and dilute pooled samples, how to dilute the QIAseq A Read1 Primer I Custom Read 1 Sequencing Primers, and then how to load a MiSeq® Cartridge. Important: If combining libraries with the same number of primers, pool equal volumes of the individual libraries at 4 nM together. If the libraries have different primer numbers, then combine the libraries at volume ratios according to their number of primers. For example, if Library A has 5000 primers at 4 nM and Library B has 600 primers at 4 nM, then combining 50 µl Library A with 6 µl Library B will result in similar coverage depth for both Libraries A and B in the same sequencing run.
9. LRMv3 Import Sample Sheet is autopopulated with information from the Protocol tab. Just fill in the highlighted yellow cells, copy this tab into a new book, then remove and delete any excess rows, and save it as a CSV file, which can then be saved where the sample sheets are saved on the Illumina® instrument. This sample sheet ensures that the Custom Read 1 Sequencing Primer is used in the sequencing run. The run conditions are 2 x 149 bp and dual indices of 10 bp each. Note: Saving CSV files on operating systems that use semicolon characters as the field separator may require you to replace all semicolons with commas using Notepad or any other text editor. Only sample sheets with commas as the field separator can be read by Illumina instruments.
QIAGEN shall not be held liable in connection with the use of this Excel tool by any user. Proper testing/validation by the respective user is recommended prior to its use. By using this tool, you expressly accept the aforementioned exclusion of QIAGEN's liability.
For up-to-date licensing information and product-specific disclaimers, see the respective QIAGEN kit handbook or user manual. QIAGEN kit handbooks and user manuals are available at www.qiagen.com or can be requested from QIAGEN Technical Services or your local distributor.
Trademarks: QIAGEN®, Sample to Insight®, QIAseq® (QIAGEN Group); BioAnalyzer®, HiSeq®, Illumina®, MiniSeq®, MiSeq® , NextSeq®, NovaSeq® (Illumina, Inc.); Excel® (Microsoft Corporation). Registered names, trademarks, etc. used in this document, even when not specifically marked as such, are not to be considered unprotected by law. © 2021 QIAGEN, all rights reserved.
Ordering www.qiagen.com/shop ¦ Technical Support support.qiagen.com ¦ Website www.qiagen.com

## Read Allocation
Formula | Reads allocated to each library (M) = coverage x primer number/106
For DNA library | DNA input required depends on the required Variant Allele Frequency (VAF).
10 ng* | Recommended coverage depends on required VAF: with four read pairs/UMI depth, 7200x coverage will give a sufficient number of UMIs and read depth for calling a 5% variant at 90% sensitivity.
20 ng* | Recommended coverage depends on required VAF: with two read pairs/UMI depth, 3640x coverage will give a sufficient number of UMIs and read depth for calling a 5% variant at 90% sensitivity.
40 ng* | Recommended coverage depends on required VAF: with four read pairs/UMI depth, 25,600x coverage will give a sufficient number of UMIs and read depth for calling a 1% variant at 90% sensitivity.
* Input of 10, 20, and 40 ng are for fresh DNA.
Instructions: | 1. Fill in cell Q18 if using a Custom Panel with the number of DNA primers.
2. Fill in cell S21 with the Final Read Allocation (M).
3. Fill in cell S33 with the Coverage (X) for DNA depending on Input and VAF.
|  | Cat. no. | DHS-001Z | DHS-002Z | DHS-003Z | DHS-005Z | DHS-101Z | DHS-102Z | DHS-103Z | DHS-104Z | DHS-105Z | DHS-3011Z | DHS-3501Z |  | Custom Panel
|  | DNA primer # | 4831 | 2929 | 5887 | 4149 | 651 | 223 | 348 | 146 | 222 | 11579 | 11311 |  |  | Fill in the DNA primer #
Per DNA library | Input | Coverage (X) | Reads allocated to each DNA library (M)
5% VAF | 10 ng | 7200 | 34.7832 | 21.0888 | 42.3864 | 29.8728 | 4.6872 | 1.6056 | 2.5056 | 1.0512 | 1.5984 | 83.3688 | 81.4392 |  | 0 |  |  | Panel
5% VAF | 20 ng | 3640 | 17.5848 | 10.6616 | 21.4287 | 15.1024 | 2.3696 | 0.81172 | 1.2667 | 0.53144 | 0.80808 | 42.1476 | 41.172 |  | 0 | Final Read Allocation (M) |  | 1
1% VAF | 40 ng | 25600 | 123.6736 | 74.9824 | 150.7072 | 106.2144 | 16.6656 | 5.7088 | 8.9088 | 3.7376 | 5.6832 | 296.4224 | 289.5616 |  | 0 |  |  | 2
3
4
The number of samples depends on the sequencing chemistry, required depth, and number of primers.
Table. Number of multiplexed samples based on a DNA panel size with XX mean coverage for DNA library per primer
|  |  | Cat. no. | DHS-001Z | DHS-002Z | DHS-003Z | DHS-005Z | DHS-101Z | DHS-102Z | DHS-103Z | DHS-104Z | DHS-105Z | DHS-3011Z | DHS-3501Z | Custom Panel | Capacity (M) | DNA library coverage (X)
Platform and Sequencing Chemistry |  |  | DNA primer # | 4831 | 2929 | 5887 | 4149 | 651 | 223 | 348 | 146 | 222 | 11579 | 11311 | 0
Instrument | Version | Read Length | Capacity (paired-ends reads) | [14 empty] | Fill in DNA coverage depending on input and VAF
MiniSeq® | Mid Output | 2 x 149 bp | 16 M | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 16
MiniSeq | High Output | 2 x 149 bp | 50 M | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 50
MiSeq | v2 Reagents | 2 x 149 bp | 30 M | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 30
MiSeq | v3 Reagents | 2 x 149 bp | 50 M | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 50
NextSeq® 500 | Mid Output | 2 x 149 bp | 260 M | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 260
NextSeq 500 | High Output | 2 x 149 bp | 800 M | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 800
HiSeq® 2500 Rapid Run | Dual Flow cell v2 | 2 x 149 bp | 1200 M | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 1200
HiSeq 3000 | 8 lanes per flow cell | 2 x 149 bp | 5 B | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 5000
HiSeq 4000 | 8 lanes per flow cell | 2 x 149 bp | 10 B | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 10000
NovaSeq® 6000 | SP (2 lanes per flow cell) | 2 x 149 bp | 1.6 B | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 1600
NovaSeq 6000 | S1 (2 lanes per flow cell) | 2 x 149 bp | 3.2 B | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 3200
NovaSeq 6000 | S2 (2 lanes per flow cell) | 2 x 149 bp | 8.2 B | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 8200
NovaSeq 6000 | S4 (4 lanes per flow cell) | 2 x 149 bp | 20 B | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | 20000

## Reagent Tracking
Reagent Tracking
Experiment: |  |  | Initials: |  | Date:
QIAseq Targeted DNA Panels |  |  |  | Cat. no.
QIAseq Targeted DNA Panel (12) |  |  |  | 333502
QIAseq Targeted DNA HC Panel (12) |  |  |  | 333512
QIAseq Targeted DNA Panel (96) |  |  |  | 333505
QIAseq Targeted DNA HC Panel (96) |  |  |  | 333515
QIAseq Targeted DNA IO Panel (12) |  |  |  | 333802
QIAseq Targeted DNA IO Panel (12) |  |  |  | 333805
QIAseq Targeted DNA Extended Panel (96) |  |  |  | 333545
QIAseq Targeted DNA Custom Panel (96) |  |  |  | 333525
QIAseq Targeted DNA Booster Panel (96) |  |  |  | 333535
QIAseq 96-Unique Dual Index Sets |  |  |  | Cat. no.
QIAseq 96-Unique Dual Index Set A |  |  |  | 333725
QIAseq 96-Unique Dual Index Set B |  |  |  | 333735
QIAseq Targeted DNA Catalog Panels |  |  |  | Cat. no. | Total number of primers* | Panel size (kb)
Human Breast Cancer Panel |  |  |  | DHS-001Z | 4831 | 370942
Human Colorectal Cancer Panel |  |  |  | DHS-002Z | 2929 | 215328
Human Myeloid Neoplasms Panel |  |  |  | DHS-003Z | 5887 | 436672
Human Lung Cancer Panel |  |  |  | DHS-005Z | 4149 | 318059
Human Actionable Solid Tumor Panel |  |  |  | DHS-101Z | 651 | 15160
Human BRCA1 and BRCA2 Panel |  |  |  | DHS-102Z | 223 | 16405
Human BRCA1 and BRCA2 Plus Panel |  |  |  | DHS-103Z | 348 | 25590
Human Pharmacogenomics Panel |  |  |  | DHS-104Z | 146 | 3313
Human HRR Panel |  |  |  | DHS-110Z | 2303 | 56485
Human Mitochondria Panel |  |  |  | DHS-105Z | 222 | 16570
Human Inherited Disease Panel |  |  |  | DHS-3011Z | 11579 | 838627
Human Comprehensive Cancer Panel |  |  |  | DHS-3501Z | 11311 | 836670
Tumor Mutational Burden Panel |  |  |  | DHS-6600Z | 19121 | 1335689
Human MSI Panel |  |  |  | DHS-7700Z | 92 | 5382
Human TMB and MSI Panel |  |  |  | DHS-8800Z | 19213 | 1314071
* The number of primers in custom, extended, and booster panels is represented by the last digits of the catalog number.
For example, a custom panel with catalog number CDHS-00100Z-1256 has 1256 primers.
Protocol
| Fragmentation, end-repair, and A-addition |  |  |  | Lot number | Date opened | Expiry date
Fragmentation Buffer, 10x
FERA Solution
FG Solution
DNase-free Water
Fragmentation Enzyme Mix
| Adapter ligation |  |  |  | Lot number | Date opened | Expiry date
Ligation Buffer, 5x
UDIN-96AA
UDIN-96BA
DNA Ligase
Ligation Solution
DNase-free water
| QIAseq Bead cleanup |  |  |  | Lot number | Date opened | Expiry date
QIAseq Beads
Ethanol
| Target enrichment |  |  |  | Lot number | Date opened | Expiry date
TEPCR Buffer, 5x
QIAseq Targeted DNA Panel
IL-F Primer
HotStarTaq DNA Polymerase
| Universal PCR |  |  |  | Lot number | Date opened | Expiry date
UPCR Buffer, 5x
HotStar Taq DNA Polymerase
DNase-free water
UDIS-96AK
UDIS-96BK

## Protocol
QIAseq Targeted DNA Panels with UDIs for Illumina Instruments with Local Run Manager
| Experiment: |  |  | Initials: |  |  | Date:
QIAseq Targeted DNA Panels with UDIs Sample Information Table
[_] | 1. Fill in Sample ID, Panel library, DNA conc. (ng/µl), and the ng per sample input amount in cell F33 for samples being tested.
[_] | 2. The dilutions will then be automatically calculated for the indicated DNA input amount in cell F33 to go directly into the
fragmentation, end-repair, and A-addition reaction.
[_] | 3. From the drop-down menus under Indices (starting at cell H35), select which well ID each of the sample will go into.
Important: Do not duplicate well ID, that is, indices. Cells in Column H will display red font if indices are duplicated, for a warning.
Also note that the well ID for each P50n and P70n will autopopulate with the correct index combination to use.
Comment: The volume amount (µl) of DNA to be added per sample and water to be added, as well as the
well ID to use, is autopopulated in the Summary Sample Table below.
Sample Summary Table
1. The table below shows how to dilute the samples appropriately for the DNA fragmentation, end-repair, and
A-addition reaction.
2. The well IDs to use in each of the index reactions are also indicated in the table below, autopopulated from the
well ID originally selected for the P70n.
ng per sample
Indices well ID
| Sample | Sample ID | Panel library | DNA conc. (ng/µl) | DNA (µl) | Water (µl) | Well ID | P50n | P70n
| 1 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 2 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 3 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 4 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 5 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 6 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 7 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 8 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 9 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 10 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 11 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 12 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 13 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 14 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 15 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 16 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 17 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 18 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 19 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 20 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 21 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 22 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 23 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 24 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 25 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 26 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 27 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 28 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 29 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 30 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 31 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 32 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 33 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 34 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 35 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 36 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 37 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 38 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 39 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 40 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 41 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 42 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 43 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 44 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 45 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 46 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 47 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 48 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 49 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 50 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 51 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 52 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 53 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 54 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 55 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 56 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 57 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 58 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 59 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 60 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 61 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 62 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 63 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 64 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 65 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 66 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 67 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 68 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 69 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 70 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 71 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 72 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 73 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 74 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 75 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 76 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 77 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 78 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 79 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 80 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 81 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 82 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 83 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 84 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 85 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 86 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 87 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 88 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 89 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 90 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 91 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 92 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 93 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 94 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 95 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| 96 |  |  |  | #DIV/0! | #DIV/0! |  | #N/A | #N/A
| Number of samples: |  |  |  | 0
Protocol: Fragmentation, End-Repair, and A-Addition
Things to do before starting:
[_] | 1. Before setting up the reaction, it is critical to accurately determine the amount of the input DNA.
[_] | 2. Set up the fragmentation, end-repair, and A-addition reaction on ice.
[_] | 3. Do not vortex any reagents or reactions.
4. Prepare the reagents required for fragmentation, end-repair, and A-addition:
[_] | a) Thaw Fragmentation Buffer, 10x; FERA Solution; and FG Solution if required at room temperature.
[_] | b) Once reagents are thawed, mix buffers thoroughly by quick vortexing to avoid any localized concentrations. Briefly
spin down vortexed reagents before use.
[_] | c) The Fragmentation Enzyme Mix should be removed from the freezer just before use and placed on ice. After use, immediately
return the enzymes to the freezer.
[_] | 5. Program thermal cyclers. For increased speed and convenience, all incubation steps of the protocol can be preprogrammed and
saved in advance.
Procedure: fragmentation, end-repair, and A-addition
[_] | 1. Enter the following cycling conditions in a PCR machine using the heated lid set to approximately 70 °C if possible. If this is not
possible, then run the cycler with the lid open for step 2 and seal the strip or plate well. When the cycler reaches step 3,
close the lid to avoid evaporation. Please spin down carefully after the run to remove any condensation.
Incubation time
| Step | Incubation temperature (°C) |  | Standard DNA | FFPE DNA | cfDNA | Mitochondria Panel
| 1 | 4 |  | 1 min | 1 min | 1 min | 1 min
| 2 | 32 |  | 24 min | 14 min | 14 min | 8 min
| 3 | 72 |  | 30 min | 30 min | 30 min | 30 min
| 4 | 4 |  | ∞ | ∞ | ∞ | ∞
[_] | 2. When the PCR cycler reaches 4°C, pause the program. It is important to follow the procedure below to achieve
optimal results.
[_] | 3. Prepare the fragmentation, end-repair, and A-addition mix for each sample as follows on ice:
Volume (µl)
|  |  |  | Standard, FFPE, or pure cfDNA |  | cfDNA contaminated with cellular DNA
| Reagents |  |  | 1 X | 0 | 1 X | 0
| DNA* |  |  | Variable |  |  | Variable | [_]
| Fragmentation Buffer, 10x |  |  | 2.5 | 0 | 2.5 | 0 | [_]
| FERA Solution |  |  | 0.75 | 0 | 0.75 | 0 | [_]
| FG Solution |  |  | 0 | 0 | 1.25 | 0 | [_]
| DNase-free water† |  |  | 1.25 | 0 | Variable | - | [_]
| Fragmentation Enzyme Mix |  |  | 5 | 0 | 5 | 0 | [_]
| Total volume |  |  | 25 | 0 | 25 | 0
* For standard DNA or cfDNA, 10–40 ng. Based on QuantiMIZE kits recommendations, use up to 250 ng of FFPE DNA. If an alternate method was used and not QuantiMIZE, then use up to 100 ng of FFPE DNA.
† This is in addition to the DNase-free water that is added from the Summary Table above.
Note: Increasing the DNA input will improve variant detection sensitivity especially for FFPE DNA samples.
[_] | 4. Mix the fragmentation, end-repair, and A-addition mix by pipetting up and down 7–8 times while keeping on ice.
Do not vortex.
[_] | 5. Add 9.5 µl of fragmentation, end-repair, and A-addition mix to the 15.5 µl of diluted DNA from the Summary Table
above.
[_] | 6. Mix by pipetting up and down 7–8 times.
[_] | 7. Seal the wells with caps (96-well plates with foil) and place in the PCR machine with the following cycling parameters:
Note: The PCR machine must be at 4°C and paused prior to adding tubes.
Incubation time
| Step | Incubation temperature (°C) |  | Standard DNA | FFPE DNA | cfDNA | Mitochondria Panel
| 1 | 4 |  | 1 min | 1 min | 1 min | 1 min
| 2 | 32 |  | 24 min | 14 min | 14 min | 8 min
| 3 | 72 |  | 30 min | 30 min | 30 min | 30 min
| 4 | 4 |  | ∞ | ∞ | ∞ | ∞
7. Immediately proceed with the "Protocol: Adapter Ligation".
Protocol: Adapter Ligation
Things to do before starting:
[_] | 1. The 25 µl product from "Protocol: Fragmentation, End-Repair, and A-addition" is the starting material for the "Protocol: Adapter
Ligation".
[_] | 2. Set up the adapter ligation reaction on ice.
[_] | 3. Do not vortex any reagents or reactions.
4. Prepare the reagents required for DNA ligation:
[_] | a) Thaw DNA Ligation Adapter; Ligation Buffer, 5x; and Ligation Solution at room temperature.
[_] | b) Once reagents are thawed, mix buffers thoroughly by quick vortexing to avoid any localized concentrations. Briefly
spin down vortexed reagents before use.
[_] | c) The DNA Ligase should be removed from the freezer just before use and placed on ice. After use, immediately
return the enzymes to the freezer.
[_] | 5. Thaw the UDIN-96AA or UDIN-96BA adapter plates on ice or store at 4°C before usage. After it is fully thawed spin down the
plate at 1000 x g for 1 min.
[_] | 6. Equilibrate the QIAseq Beads to room temperature for 20–30 min before use. Ensure that the QIAseq Beads are thoroughly mixed at
all times.
[_] | 7. Prepare fresh 80% ethanol.
Procedure: adapter ligation
[_] | 1. Prepare the adapter ligation mix for each sample as follows on ice:
Volume (µl)
[5 empty] | Standard and FFPE DNA |  | cfDNA
| Reagents |  |  |  | 1 X | 0 | 1 X | 0
| Fragmentation, end-repair, and A-addition reaction (already in tube) |  |  |  | 25 | - | 25 | - | [_]
| Ligation Buffer, 5x |  |  |  | 10 | 0 | 10 | 0 | [_]
| UDIN-96AA or UDIN-96BA Adapter Plate |  |  |  | 2.8 | - | 0.5 | - | [_]
| DNA Ligase |  |  |  | 5 | 0 | 5 | 0 | [_]
| Ligation Solution |  |  |  | 7.2 | 0 | 7.2 | 0 | [_]
| DNase-free water |  |  |  | 0 | 0 | 2.3 | 0 | [_]
| Total volume |  |  |  | 22.2 | 0 | 24.5 | 0
Note: The Ligation Solution is extremely viscous. Pipette slowly.
[_] | 2. Thoroughly mix by pipetting up and down at least 10–12 times. Do not vortex.
Note: The UDIN-96#A: Set A or B adapter plates used in the ligation must be paired with the UDIS-96#K Primer Plate: Set A or B,
respectively, in the universal PCR step.
[_] | 3. Add 22.2 µl ligation master mix for standard and FFPE DNA to the 25 µl sample from the
fragmented/end-repair/A-addition reaction or add 24.5 µl ligation master mix for cfDNA
to the 25 µl sample from the fragmented/end-repair/A-addition reaction.
[_] | 4. Add either 2.8 µl for standard and FFPE DNA or 0.5 µl for cfDNA of the appropriate P7 index to the individual samples.
Note: The UDIN-96#A adapters are sealed in a 96-well plate with pierceable aluminum heat sealing film. Puncture the film with a
standard 200 µl pipet tip to withdraw the appropriate adapter and adapter volume.
Remember to use a unique P7 index for each individual sample.
[_] | 5. Incubate the 50 µl ligation reaction at 20°C for exactly 15 min using a PCR cycler with the lid open.
4°C for 1 min; 20°C for 15 min; 4°C for ∞
DO NOT USE A HEATED LID
6. Once the run has finished add the appropriate volume of water to the adapter-ligated DNA sample and transfer the reaction to
a clean 1.5 ml Lo-Bind Tube.
[_] | a) Add 50 µl of water to bring the reaction 100 µl for standard DNA and FFPE samples.
[_] | b) Add 30 µl of water to bring the reaction 80 µl for cfDNA samples.
7. Vortex the QIAseq Beads prior to:
[_] | a) Add 100 µl of the QIAseq Beads to the 100 µl standard DNA and FFPE reactions.
[_] | b) Add 112 µl of the QIAseq Beads to the cfDNA reactions.
Mix well by pipetting 10 times.
[_] | 8. Incubate for 5 min at room temperature.
[_] | 9. Place the reaction on a magnet to separate the beads from the supernatant. It should be clear after 2 min (for tubes) or
10 min (for plates). Carefully remove and discard the supernatant as the beads contain the target.
[_] | 10. Completely remove all residual supernatant.
[_] | 11. Wash the beads with 200 µl of freshly prepared 80% ethanol.
[_] | 12. Remove the 80% ethanol wash being careful not to disturb the beads.
[_] | 13. Wash the beads with 200 µl of freshly prepared 80% ethanol again.
[_] | 14. Remove the 80% ethanol wash and the residual ethanol being careful not to disturb the beads.
[_] | 15. Air dry the beads for 10 min. Check by eye to ensure that the beads are dried.
[_] | 16. Elute the DNA targets into 52 µl of water by pipetting up and down.
[_] | 17. Incubate for 5 min at room temperature prior to separating the supernatant from the beads by placing the reaction on a
magnet until the solution is clear which should take approximately 2 min.
[_] | 18. Transfer 50 µl of the supernatant to a clean tube or PCR tube.
19. Vortex the QIAseq Beads prior to:
[_] | a) Add 50 µl of the QIAseq Beads to the 50 µl standard DNA and FFPE reactions.
[_] | b) Add 35 µl of the QIAseq Beads to the Human Mitochondrial Panel Samples.
[_] | c) Add 70 µl of the QIAseq Beads to the cfDNA reactions.
Mix well by pipetting 10 times.
[_] | 20. Incubate for 5 min at room temperature.
[_] | 21. Place the reaction on a magnet to separate the beads from the supernatant. It should be clear after 5 min. Carefully
remove and discard the supernatant as the beads contain the target.
[_] | 22. Completely remove all residual supernatant.
[_] | 23. Wash the beads with 200 µl of freshly prepared 80% ethanol.
[_] | 24. Remove the 80% ethanol wash being careful not to disturb the beads.
[_] | 25. Wash the beads with 200 µl of freshly prepared 80% ethanol again.
[_] | 26. Remove the 80% ethanol wash and the residual ethanol being careful not to disturb the beads.
[_] | 27. Air dry the beads for 15 min. Check by eye to ensure that the beads are dried.
[_] | 28. Elute the DNA targets into 12 µl of water by pipetting up and down.
[_] | 29. Incubate for 5 min at room temperature prior to separating the supernatant from the beads by placing the reaction on a
magnet until the solution is clear which should take approximately 2 min.
[_] | 30. Transfer 9.4 µl of the supernatant to a clean tube or PCR tube.
[_] | 31. Upon completion, place the reactions on ice and proceed with the "Protocol: Target Enrichment"; alternatively, the
samples can be stored at −30 to −15°C in a constant-temperature freezer. This is a safe stopping point.
Protocol: Target Enrichment
Things to do before starting:
[_] | 1. The 9.4 µl product from "Protocol: Adapter Ligation" is the starting material for the "Protocol: Target Enrichment".
[_] | 2. Set up the target enrichment reaction on ice.
[_] | 3. Do not vortex any reagents or reactions.
4. Prepare the reagents required for target enrichment:
[_] | a) Thaw TEPCR Buffer, 5x; QIAseq Targeted DNA Panel; and IL-Forward Primer at room temperature.
[_] | b) Once reagents are thawed, mix buffers thoroughly by quick vortexing to avoid any localized concentrations. Briefly
spin down vortexed reagents before use.
[_] | c) The HotStarTaq should be removed from the freezer just before use and placed on ice. After use, immediately
return the enzymes to the freezer.
[_] | 5. Equilibrate the QIAseq Beads to room temperature for 20–30 min before use. Ensure that the QIAseq Beads are thoroughly mixed at
all times.
[_] | 6. Prepare fresh 80% ethanol.
Procedure: Target enrichment
[_] | 1. Prepare the following target enrichment mix:
Volume (µl)
|  | Reagents |  |  |  | 1 X | 0
|  | TEPCR Buffer, 5x |  |  |  | 4 | 0 | [_]
|  | QIAseq Targeted DNA Panel |  |  |  | 5 | 0 | [_]
|  | IL-F Primer |  |  |  | 0.8 | 0 | [_]
|  | Adapter-ligated DNA from “Protocol: Adapter Ligation” |  |  |  | 9.4 | - | [_]
|  | HotStarTaq DNA Polymerase |  |  |  | 0.8 | 0 | [_]
|  | Total volume |  |  |  | 20 | 0
[_] | 2. Add 10.6 µl of the target enrichment mix to the 9.4 µl of the adapter-ligated DNA from “Protocol: Adapter Ligation”.
[_] | 3. Seal the wells with caps and place in the PCR machine with the following cycling parameters:
95°C for 13 min, 98°C for 2 min (1 cycle)
98°C for 15 s, 68 °C for 10 min (8 cycles)
72°C for 5 min (1 cycle)
4°C for 5 min (1 cycle)
4°C for ∞
NOTE: If the number of primers in the panel is ≥ 1500 per tube use the below conditions instead:
95°C for 13 min, 98°C for 2 min (1 cycle)
98°C for 15 s, 65°C for 15 min (6 cycles)
72°C for 5 min (1 cycle)
4°C for 5 min (1 cycle)
4°C for ∞
4. Once the run has finished, add the appropriate volume of water to the targeted enriched DNA sample and transfer the
reaction to a clean 1.5 ml Lo-Bind Tube.
[_] | a) Add 80 µl of water to bring the reaction 100 µl for standard DNA and FFPE samples.
[_] | b) Add 70 µl of water to bring the reaction 90 µl for cfDNA samples.
5. Vortex the QIAseq Beads prior to:
[_] | a) Add 100 µl of the QIAseq Beads to the 100 µl standard DNA and FFPE reactions.
[_] | b) Add 70 µl of the QIAseq Beads to the Human Mitochondrial Panel Samples.
[_] | c) Add 108 µl of the QIAseq Beads to the cfDNA reactions.
Mix well by pipetting 10 times.
[_] | 6. Incubate for 5 min at room temperature.
[_] | 7. Place the reaction on a magnet to separate the beads from the supernatant. It should be clear after 2 min (for tubes) or
10 min (for plates). Carefully remove and discard the supernatant as the beads contain the target.
[_] | 8. Completely remove all residual supernatant.
[_] | 9. Wash the beads with 200 µl of freshly prepared 80% ethanol.
[_] | 10. Remove the 80% ethanol wash being careful not to disturb the beads.
[_] | 11. Wash the beads with 200 µl of freshly prepared 80% ethanol again.
[_] | 12. Remove the 80% ethanol wash and the residual Ethanol being careful not to disturb the beads.
[_] | 13. Air dry the beads for 10 min. Check by eye to ensure that the beads are dried.
[_] | 14. Elute the DNA targets into 16 µl of water by pipetting up and down.
[_] | 15. Incubate for 5 min at room temperature prior to separating the supernatant from the beads by placing the reaction on a
magnet until the solution is clear which should take approximately 2 min.
[_] | 16. Transfer 13.4 µl of the supernatant to a clean tube or PCR tube.
[_] | 17. Proceed with the "Protocol: Universal PCR"; alternatively, the samples can be stored at −30 to −15°C in a constant-temperature
freezer. This is a safe stopping point.
Protocol: Universal PCR
Things to do before starting:
[_] | 1. The 13.4 µl product from "Protocol: Target Enrichment" is the starting material for the "Protocol: Universal PCR".
[_] | 2. Set up the universal PCR reaction on ice.
[_] | 3. Do not vortex any reagents or reactions.
4. Prepare the reagents required for universal PCR:
[_] | a) Thaw UPCR Buffer, 5x, at room temperature.
[_] | b) Once reagents are thawed, mix buffers thoroughly by quick vortexing to avoid any localized concentrations. Briefly
spin down vortexed reagents before use.
[_] | c) The HotStarTaq should be removed from the freezer just before use and placed on ice. After use, immediately
return the enzymes to the freezer.
[_] | 5. Bring the UDIS-96#K Index Primer Plate A or B containing pre-dispense, dried index primers, and the universal PCR primers to
room temperature.
[_] | 6. Equilibrate the QIAseq Beads to room temperature for 20–30 min before use. Ensure that the QIAseq Beads are thoroughly mixed at
all times.
[_] | 7. Prepare fresh 80% ethanol.
Procedure: universal PCR
[_] | 1. Prepare the universal PCR mix as follows:
Volume (µl)
|  | Reagents |  |  | 1 X | 0
|  | UPCR Buffer, 5x |  |  | 4 | 0 | [_]
|  | Sample from "Protocol: Target Enrichment” |  |  | 13.4 | - | [_]
|  | HotStarTaq DNA Polymerase |  |  | 1 | 0 | [_]
|  | DNase-free Water |  |  | 1.6 | 0 | [_]
|  | Total Volume |  |  | 20 | 0
[_] | 2. Mix the universal PCR reaction mix by pipetting up and down 10–12 times. Briefly centrifuge.
[_] | 3. Add 6.6 µl of the universal PCR reaction mix and 13.4 µl of the target-enriched DNA from "Protocol: Target Enrichment"
from above directly in the UDIS-96#K plate. Each sample needs a unique i5 Index. The total volume is 20 µl.
Note: The UDIS-96#K Index Primer Plate A or B is dried down in a 96-well plate. Please see the format in the Index Plates tab.
The UDIS-96#K Index Primers are sealed in a 96-well plate with pierceable aluminum heat sealing film. Puncture the film with a
standard 200 µl pipet tip to withdraw the appropriate adapter and adapter volume.
[_] | 4. Mix by pipetting up and down 10–12 times. Briefly centrifuge.
[_] | 5. Seal the wells with caps and place in the PCR machine with the following cycling parameters:
95°C for 13 min, 98°C for 2 min (1 cycle)
98°C for 15 s, 60°C for 2 min (Y cycles*)
72°C for 5 min (1 cycle)
4°C for 5 min (1 cycle)
4°C for ∞
* Where "Y cycles" is based on the following table:
Cycle number
|  |  | Primers per pool | Standard DNA | cfDNA | FFPE DNA
|  |  | 6–24 | 26 | 28 | 30
|  |  | 25–96 | 24 | 26 | 28
|  |  | 97–288 | 22 | 24 | 26
|  |  | 289–1056 | 21 | 23 | 25
|  |  | 1057–1499 | 20 | 22 | 24
|  |  | 1500–3072 | 21 | 23 | 25
|  |  | 3073–4999 | 20 | 22 | 24
|  |  | 5000–12,000 | 19 | 21 | 23
|  |  | ≥ 12,001 | 18 | 20 | 22
6. After the cycling is complete, place the reactions on ice and add the appropriate volume of water to the universal PCR samples and
transfer the reaction to a clean 1.5 ml Lo-Bind Tube.
[_] | a) Add 80 µl of water to bring the reaction 100 µl for standard DNA and FFPE samples.
[_] | b) Add 70 µl of water to bring the reaction 90 µl for cfDNA samples.
7. Vortex the QIAseq Beads prior to:
[_] | a) Add 100 µl of the QIAseq Beads to the 100 µl standard DNA and FFPE reactions.
[_] | b) Add 70 µl of the QIAseq Beads to the Human Mitochondrial Panel Samples.
[_] | c) Add 108 µl of the QIAseq Beads to the cfDNA reactions.
Mix well by pipetting 10 times.
[_] | 8. Incubate for 5 min at room temperature.
[_] | 9. Place the reaction on a magnet to separate the beads from the supernatant. It should be clear after 2 min (for tubes) or
10 min (for plates). Carefully remove and discard the supernatant as the beads contain the target.
[_] | 10. Completely remove all residual supernatant.
[_] | 11. Wash the beads with 200 µl of freshly prepared 80% ethanol.
[_] | 12. Remove the 80% ethanol wash being careful not to disturb the beads.
[_] | 13. Wash the beads with 200 µl of freshly prepared 80% ethanol again.
[_] | 14. Remove the 80% ethanol wash and the residual ethanol being careful not to disturb the beads.
[_] | 14. Air dry the beads for 10 min. Check by eye to ensure that the beads are dried.
[_] | 16. Elute the DNA targets into 30 µl of water by pipetting up and down.
[_] | 17. Incubate for 5 min at room temperature prior to separating the supernatant from the beads by placing the reaction on a magnet
until the solution is clear which should take approximately 2 min.
[_] | 18. Transfer 28 µl of the supernatant to a clean tube or PCR tube.
[_] | 19. Proceed to Quantitate the Libraries either by QIAseq Library Quant Array for Illumina or BioAnalyzer® Measurements; alternatively,
the samples can be stored at –30 to –15°C in a constant-temperature freezer. This is a safe stopping point.

## Autopopulate UDI's (hidden)
| i7 Index Name | i7 Bases for Entry on Sample Sheet |  | i5 Index Name | i5 Bases for Entry on DNA Sample Sheet | Platform and Sequencing Chemistry
PlateA_A01 | NQDIB001 | ACAACGTTCA | PlateA_A01 | SQDIB001 | ATGGCCGACT | Instrument | Version | Instrument_Reagent | Capacity (paired-ends reads) | Capacity (M)
PlateA_B01 | NQDIB002 | CAAGTCTGGT | PlateA_B01 | SQDIB002 | CGATGAGCAC | MiniSeq | Mid Output | MiniSeq_Mid Output | 16 M | 16
PlateA_C01 | NQDIB003 | GTTCGCCAGT | PlateA_C01 | SQDIB003 | GATAAGTCGA | MiniSeq | High Output | MiniSeq_High Output | 50 M | 50
PlateA_D01 | NQDIB004 | TGCCTAACGC | PlateA_D01 | SQDIB004 | TCACGCCTTG | MiSeq | v2 Reagents | MiSeq_v2 Reagents | 30 M | 30
PlateA_E01 | NQDIB005 | AGGCCGATAA | PlateA_E01 | SQDIB005 | AGGAACACAA | MiSeq | v3 Reagents | MiSeq_v3 Reagents | 50 M | 50
PlateA_F01 | NQDIB006 | CTTATACCTC | PlateA_F01 | SQDIB006 | CTCAGTAGGC | NextSeq 500 | Mid Output | NextSeq 500_Mid Output | 260 M | 260
PlateA_G01 | NQDIB007 | GAATCCTTGA | PlateA_G01 | SQDIB007 | GAAGTGCCTG | NextSeq 500 | High Output | NextSeq 500_High Output | 800 M | 800
PlateA_H01 | NQDIB008 | TCTCGGTTCG | PlateA_H01 | SQDIB008 | TCTCTCGCCT | HiSeq 2500 Rapid Run | Dual Flow cell v2 | HiSeq 2500 Rapid Run _Dual Flow cell v2 | 1200 M | 1200
PlateA_A02 | NQDIB009 | AACTTGGCTC | PlateA_A02 | SQDIB009 | AGGCACCTTC | HiSeq 3000 | 8 lanes per flow cell | HiSeq 3000_8 lanes per flow cell | 5 B | 5000
PlateA_B02 | NQDIB010 | CTACGGCCTT | PlateA_B02 | SQDIB010 | CTGTTGGTAA | HiSeq 4000 | 8 lanes per flow cell | HiSeq 4000_8 lanes per flow cell | 10 B | 10000
PlateA_C02 | NQDIB011 | GCTTCTCTAA | PlateA_C02 | SQDIB011 | GCTGGTACCT | NovaSeq 6000 | SP (2 lanes per flow cell) | NovaSeq 6000_SP (2 lanes per flow cell) | 1.6 B | 1600
PlateA_D02 | NQDIB012 | TGGTCTTAGA | PlateA_D02 | SQDIB012 | TAAGGAGCGG | NovaSeq 6000 | S1 (2 lanes per flow cell) | NovaSeq 6000_S1 (2 lanes per flow cell) | 3.2 B | 3200
PlateA_E02 | NQDIB013 | AGTGGTTACA | PlateA_E02 | SQDIB013 | AATCGCTCCA | NovaSeq 6000 | S2 (2 lanes per flow cell) | NovaSeq 6000_S2 (2 lanes per flow cell) | 8.2 B | 8200
PlateA_F02 | NQDIB014 | CTTGTGTCGG | PlateA_F02 | SQDIB014 | CTCCTAATTG | NovaSeq 6000 | S4 (4 lanes per flow cell) | NovaSeq 6000_S4 (4 lanes per flow cell) | 20 B | 20000
PlateA_G02 | NQDIB015 | GCCATCAGAG | PlateA_G02 | SQDIB015 | GCCTCATAAT
PlateA_H02 | NQDIB016 | TAACAGGCCG | PlateA_H02 | SQDIB016 | TGTATTGAGC
PlateA_A03 | NQDIB017 | AAGCTATGCA | PlateA_A03 | SQDIB017 | AGCCATAACA
PlateA_B03 | NQDIB018 | CGAGAAGGTT | PlateA_B03 | SQDIB018 | CCACAAGTGG
PlateA_C03 | NQDIB019 | GTGATCTCTT | PlateA_C03 | SQDIB019 | GTTATCACAC
PlateA_D03 | NQDIB020 | TCCTTCAGGC | PlateA_D03 | SQDIB020 | TACCGTTCTT
PlateA_E03 | NQDIB021 | AAGGCACAAT | PlateA_E03 | SQDIB021 | AGGCGTTAGG
PlateA_F03 | NQDIB022 | CGGTAGAGGA | PlateA_F03 | SQDIB022 | CCGTAACGTC
PlateA_G03 | NQDIB023 | GTTCATGGTA | PlateA_G03 | SQDIB023 | GTAATAGCCA
PlateA_H03 | NQDIB024 | TCTGCCAATG | PlateA_H03 | SQDIB024 | TAGCGCCGAT
PlateA_A04 | NQDIB025 | AATAGCAGTG | PlateA_A04 | SQDIB025 | CATTCTTGGA
PlateA_B04 | NQDIB026 | ACCTACCATT | PlateA_B04 | SQDIB026 | ATGCAAGGTT
PlateA_C04 | NQDIB027 | CATAGGTATC | PlateA_C04 | SQDIB027 | CGCCAGACAA
PlateA_D04 | NQDIB028 | GTACCTAGTG | PlateA_D04 | SQDIB028 | GAAGGTTGGC
PlateA_E04 | NQDIB029 | TGAACGAGCT | PlateA_E04 | SQDIB029 | TCGCATCACG
PlateA_F04 | NQDIB030 | AAGACTGACA | PlateA_F04 | SQDIB030 | CCGGTCATGA
PlateA_G04 | NQDIB031 | ACTGTTCATC | PlateA_G04 | SQDIB031 | ATTCACAAGC
PlateA_H04 | NQDIB032 | CGCCGATTGT | PlateA_H04 | SQDIB032 | CAACCTGTAA
PlateA_A05 | NQDIB033 | GAACTCAATC | PlateA_A05 | SQDIB033 | GCCAGTCGTT
PlateA_B05 | NQDIB034 | TTGGCATTAC | PlateA_B05 | SQDIB034 | TGCCTTGTCG
PlateA_C05 | NQDIB035 | AGCGCAACGA | PlateA_C05 | SQDIB035 | CTATCCGCTG
PlateA_D05 | NQDIB036 | ATACTCACCT | PlateA_D05 | SQDIB036 | AATGCCGGAA
PlateA_E05 | NQDIB037 | CCATTATCGA | PlateA_E05 | SQDIB037 | CGGTTATCCG
PlateA_F05 | NQDIB038 | GAAGAGACGC | PlateA_F05 | SQDIB038 | GCGGAAGAGT
PlateA_G05 | NQDIB039 | TGCAGGAGAC | PlateA_G05 | SQDIB039 | TTGGTTAGTC
PlateA_H05 | NQDIB040 | AATGAAGCTC | PlateA_H05 | SQDIB040 | TTCAGTGTGA
PlateA_A06 | NQDIB041 | ATGTAGGCCT | PlateA_A06 | SQDIB041 | AGAATTCTGG
PlateA_B06 | NQDIB042 | CGGTTCCACA | PlateA_B06 | SQDIB042 | CATTGACTCT
PlateA_C06 | NQDIB043 | GCTTAATACG | PlateA_C06 | SQDIB043 | GCGGCTTCAA
PlateA_D06 | NQDIB044 | TAACCACTGG | PlateA_D06 | SQDIB044 | TTATGGTCTC
PlateA_E06 | NQDIB045 | ACTCGAACGC | PlateA_E06 | SQDIB045 | CGTAACCAGG
PlateA_F06 | NQDIB046 | AACCGGAAGG | PlateA_F06 | SQDIB046 | AGCTCAGATA
PlateA_G06 | NQDIB047 | CCGTCTTGTG | PlateA_G06 | SQDIB047 | CCGGTGTTAC
PlateA_H06 | NQDIB048 | GTGTGTAAGC | PlateA_H06 | SQDIB048 | GACCTAACCT
PlateA_A07 | NQDIB049 | TGGACATCCT | PlateA_A07 | SQDIB049 | TTGTAGAAGG
PlateA_B07 | NQDIB050 | ACATAAGGTG | PlateA_B07 | SQDIB050 | CCTAGCACTA
PlateA_C07 | NQDIB051 | ACAGCCGCTT | PlateA_C07 | SQDIB051 | ATCGTGTTCT
PlateA_D07 | NQDIB052 | CTCACAGGAA | PlateA_D07 | SQDIB052 | CCAACTTATC
PlateA_E07 | NQDIB053 | GAACTGTACT | PlateA_E07 | SQDIB053 | GAAGCCAAGG
PlateA_F07 | NQDIB054 | TGAGGCTGTA | PlateA_F07 | SQDIB054 | TGGAGTTCAA
PlateA_G07 | NQDIB055 | CCAATAGAAC | PlateA_G07 | SQDIB055 | CTTCAATCCT
PlateA_H07 | NQDIB056 | ACCGGTATAT | PlateA_H07 | SQDIB056 | ATCTTGCGTG
PlateA_A08 | NQDIB057 | CATTCCGAGG | PlateA_A08 | SQDIB057 | CGTCTAAGGT
PlateA_B08 | NQDIB058 | GTTCCAGAAC | PlateA_B08 | SQDIB058 | GAGGTGAACA
PlateA_C08 | NQDIB059 | TGGTGAATCT | PlateA_C08 | SQDIB059 | TCAGAACTAC
PlateA_D08 | NQDIB060 | ATCTGACCGA | PlateA_D08 | SQDIB060 | CGGATATTGA
PlateA_E08 | NQDIB061 | AGCGAGAGTG | PlateA_E08 | SQDIB061 | AGGAGTAGAT
PlateA_F08 | NQDIB062 | CTGGACCAAC | PlateA_F08 | SQDIB062 | CCGCCGAATA
PlateA_G08 | NQDIB063 | GCTTCGAGCT | PlateA_G08 | SQDIB063 | GAGTCTATAC
PlateA_H08 | NQDIB064 | TAGAACCTCT | PlateA_H08 | SQDIB064 | TTATTACCGG
PlateA_A09 | NQDIB065 | TTCGAGTCAT | PlateA_A09 | SQDIB065 | CGCTCGTTAG
PlateA_B09 | NQDIB066 | AGGATTGTTC | PlateA_B09 | SQDIB066 | AACAACGCTG
PlateA_C09 | NQDIB067 | CTCCTTGCCA | PlateA_C09 | SQDIB067 | CGCGGCTATT
PlateA_D09 | NQDIB068 | GCCAATATTC | PlateA_D09 | SQDIB068 | GCTCGACACA
PlateA_E09 | NQDIB069 | TAGGTTCCGG | PlateA_E09 | SQDIB069 | TTCTTCCAAC
PlateA_F09 | NQDIB070 | CCGAACAAGT | PlateA_F09 | SQDIB070 | TTGGCGGTTG
PlateA_G09 | NQDIB071 | ATTGGACTTG | PlateA_G09 | SQDIB071 | AACAGGCAAT
PlateA_H09 | NQDIB072 | CCTTGCGGTT | PlateA_H09 | SQDIB072 | CAGAATGGCG
PlateA_A10 | NQDIB073 | GAGTCAACGT | PlateA_A10 | SQDIB073 | GTTGAGATTC
PlateA_B10 | NQDIB074 | TGTTAAGTGG | PlateA_B10 | SQDIB074 | TGTGTGCGGA
PlateA_C10 | NQDIB075 | AGGAACTGCT | PlateA_C10 | SQDIB075 | GTTCGGCGAA
PlateA_D10 | NQDIB076 | ACGAAGGCGA | PlateA_D10 | SQDIB076 | AGCTGTATTG
PlateA_E10 | NQDIB077 | CGCAGTCCTA | PlateA_E10 | SQDIB077 | CAGCGGATGA
PlateA_F10 | NQDIB078 | TTCGCTCGGA | PlateA_F10 | SQDIB078 | GTCCTTGGAT
PlateA_G10 | NQDIB079 | AACAACCGAA | PlateA_G10 | SQDIB079 | TCTAGATGCT
PlateA_H10 | NQDIB080 | TTCCTCCTGT | PlateA_H10 | SQDIB080 | CGAGCCACAT
PlateA_A11 | NQDIB081 | TTAATGGAGG | PlateA_A11 | SQDIB081 | ATGGAATGGA
PlateA_B11 | NQDIB082 | AACCGCGACT | PlateA_B11 | SQDIB082 | CATTCCTCAC
PlateA_C11 | NQDIB083 | CCTGGATGAG | PlateA_C11 | SQDIB083 | GCATAGGAAG
PlateA_D11 | NQDIB084 | TTCAACCACA | PlateA_D11 | SQDIB084 | TGTTCGTGTT
PlateA_E11 | NQDIB085 | CCACGCATAA | PlateA_E11 | SQDIB085 | TAAGACCGTT
PlateA_F11 | NQDIB086 | ATACATTCGC | PlateA_F11 | SQDIB086 | ATGGTACCAG
PlateA_G11 | NQDIB087 | CGAGCTTGAC | PlateA_G11 | SQDIB087 | CCGACAGCTT
PlateA_H11 | NQDIB088 | TCCAACTCTA | PlateA_H11 | SQDIB088 | GACGATATGA
PlateA_A12 | NQDIB089 | AGATCATCAG | PlateA_A12 | SQDIB089 | TTGTACTCCA
PlateA_B12 | NQDIB090 | AACACCTAGT | PlateA_B12 | SQDIB090 | GTGCACATAA
PlateA_C12 | NQDIB091 | CCGCTAACAG | PlateA_C12 | SQDIB091 | AGGACAAGTA
PlateA_D12 | NQDIB092 | TTGGTGCGAT | PlateA_D12 | SQDIB092 | CCGATTCGAG
PlateA_E12 | NQDIB093 | ACCAAGTAAG | PlateA_E12 | SQDIB093 | GTAGGAACTT
PlateA_F12 | NQDIB094 | CGCATTAAGG | PlateA_F12 | SQDIB094 | TACACTACGA
PlateA_G12 | NQDIB095 | CTAGGCGAGA | PlateA_G12 | SQDIB095 | ATGACCTTGA
PlateA_H12 | NQDIB096 | TCTCTGAAGA | PlateA_H12 | SQDIB096 | CTACGTGACG
PlateB_A01 | NQDIB097 | TACCGGTGGT | PlateB_A01 | SQDIB097 | AACAATCAGG
PlateB_B01 | NQDIB098 | AGGTGTTACG | PlateB_B01 | SQDIB098 | CTGGTGTGCA
PlateB_C01 | NQDIB099 | ACAGACCGAC | PlateB_C01 | SQDIB099 | GCATATCCTT
PlateB_D01 | NQDIB100 | CGAATACGTA | PlateB_D01 | SQDIB100 | TGTCCTGTAC
PlateB_E01 | NQDIB101 | TAGCATCGAT | PlateB_E01 | SQDIB101 | AGAACGTCGC
PlateB_F01 | NQDIB102 | CCATGAGTCG | PlateB_F01 | SQDIB102 | CACGGACTAG
PlateB_G01 | NQDIB103 | ACTAACATGC | PlateB_G01 | SQDIB103 | GTTGAACACT
PlateB_H01 | NQDIB104 | ACACTCTCTA | PlateB_H01 | SQDIB104 | TCGCGTGGTA
PlateB_A02 | NQDIB105 | GCTCTTGCCT | PlateB_A02 | SQDIB105 | AGCCACTATG
PlateB_B02 | NQDIB106 | AATCTTGAGG | PlateB_B02 | SQDIB106 | CCACCTACCA
PlateB_C02 | NQDIB107 | CTTAACGGTC | PlateB_C02 | SQDIB107 | GTTCCGGTGT
PlateB_D02 | NQDIB108 | TTGTGACCAA | PlateB_D02 | SQDIB108 | TAGGTCTGAC
PlateB_E02 | NQDIB109 | TCACACACCT | PlateB_E02 | SQDIB109 | AGGAAGCATT
PlateB_F02 | NQDIB110 | CTGCAATTAG | PlateB_F02 | SQDIB110 | CCTTAGTTGG
PlateB_G02 | NQDIB111 | CTCCTTACTC | PlateB_G02 | SQDIB111 | GTCCTATTCA
PlateB_H02 | NQDIB112 | GCAACGCAGA | PlateB_H02 | SQDIB112 | TAAGATGGAC
PlateB_A03 | NQDIB113 | CCTTACCAAT | PlateB_A03 | SQDIB113 | AGGCCATGGT
PlateB_B03 | NQDIB114 | TTAATCCTCG | PlateB_B03 | SQDIB114 | CATTGGCCAA
PlateB_C03 | NQDIB115 | TTCCGAGTTC | PlateB_C03 | SQDIB115 | GCTATGAATC
PlateB_D03 | NQDIB116 | CTCGAGAGGA | PlateB_D03 | SQDIB116 | TTGGTCCTCG
PlateB_E03 | NQDIB117 | TGTTGGCTGT | PlateB_E03 | SQDIB117 | AGCGACATAC
PlateB_F03 | NQDIB118 | CGTATCTGCG | PlateB_F03 | SQDIB118 | CAAGTAGTCT
PlateB_G03 | NQDIB119 | CCATAGTATC | PlateB_G03 | SQDIB119 | GTCAAGAAGA
PlateB_H03 | NQDIB120 | TGGACAGTAA | PlateB_H03 | SQDIB120 | TCCTGTTATG
PlateB_A04 | NQDIB121 | GTACCTTGTT | PlateB_A04 | SQDIB121 | AAGTGCGATA
PlateB_B04 | NQDIB122 | GAGTGCCTCT | PlateB_B04 | SQDIB122 | AGGCTACACG
PlateB_C04 | NQDIB123 | TAAGTAGCGG | PlateB_C04 | SQDIB123 | CTATATCGGC
PlateB_D04 | NQDIB124 | CGTGGTGTTC | PlateB_D04 | SQDIB124 | GCTAAGGTAA
PlateB_E04 | NQDIB125 | CATTCCTGAA | PlateB_E04 | SQDIB125 | TAACCTGGTT
PlateB_F04 | NQDIB126 | AAGATGCATG | PlateB_F04 | SQDIB126 | AGTTGGTCTA
PlateB_G04 | NQDIB127 | CCTTGGAGCT | PlateB_G04 | SQDIB127 | ATGCAGCTGG
PlateB_H04 | NQDIB128 | ACCGGAACAG | PlateB_H04 | SQDIB128 | CGTTGCCTTC
PlateB_A05 | NQDIB129 | GAATGGAAGC | PlateB_A05 | SQDIB129 | GCGTGGAGAA
PlateB_B05 | NQDIB130 | GTTCTCCATA | PlateB_B05 | SQDIB130 | TACGCCTCCT
PlateB_C05 | NQDIB131 | GTCACTATGT | PlateB_C05 | SQDIB131 | AATTCGGTAG
PlateB_D05 | NQDIB132 | TGGTAGAACT | PlateB_D05 | SQDIB132 | ATTGTCGAAC
PlateB_E05 | NQDIB133 | ACGCCTATGG | PlateB_E05 | SQDIB133 | CAACCTTGCG
PlateB_F05 | NQDIB134 | AATCCGTTAC | PlateB_F05 | SQDIB134 | GCACTGCGTA
PlateB_G05 | NQDIB135 | GTTGAGGCTA | PlateB_G05 | SQDIB135 | TGCTAGTAGT
PlateB_H05 | NQDIB136 | TATCAACTGG | PlateB_H05 | SQDIB136 | AAGTCACGGA
PlateB_A06 | NQDIB137 | AAGAGGAGAT | PlateB_A06 | SQDIB137 | AGCGATTGAA
PlateB_B06 | NQDIB138 | GTCTTCTCGG | PlateB_B06 | SQDIB138 | CTACCTCTCT
PlateB_C06 | NQDIB139 | GAAGCCACTC | PlateB_C06 | SQDIB139 | GACAACTGTC
PlateB_D06 | NQDIB140 | GTAGGACACA | PlateB_D06 | SQDIB140 | TCCATTGCGG
PlateB_E06 | NQDIB141 | CTCCTCGTAT | PlateB_E06 | SQDIB141 | AGCCTCGCAA
PlateB_F06 | NQDIB142 | CCACATGATT | PlateB_F06 | SQDIB142 | AATACAGGCT
PlateB_G06 | NQDIB143 | AGACGGTTGG | PlateB_G06 | SQDIB143 | CGGACCGTTA
PlateB_H06 | NQDIB144 | CTAGGTTGAC | PlateB_H06 | SQDIB144 | GCGCTTATGC
PlateB_A07 | NQDIB145 | AAGCGTACCA | PlateB_A07 | SQDIB145 | TTAACACGAG
PlateB_B07 | NQDIB146 | TCATGTTGGT | PlateB_B07 | SQDIB146 | CGCCTCTAGA
PlateB_C07 | NQDIB147 | TTGGAATGGT | PlateB_C07 | SQDIB147 | AATCGACCTT
PlateB_D07 | NQDIB148 | GTGTATGTTG | PlateB_D07 | SQDIB148 | CCGCAATAAC
PlateB_E07 | NQDIB149 | TCCTGTCAAC | PlateB_E07 | SQDIB149 | GTTCCAACGA
PlateB_F07 | NQDIB150 | TAATCAGGCA | PlateB_F07 | SQDIB150 | TGTTAGACCG
PlateB_G07 | NQDIB151 | GTAGTGGATT | PlateB_G07 | SQDIB151 | AACCTCATAG
PlateB_H07 | NQDIB152 | AATTGCGCAT | PlateB_H07 | SQDIB152 | ATGAATCCAC
PlateB_A08 | NQDIB153 | GACAATAACG | PlateB_A08 | SQDIB153 | CGGCTTAATT
PlateB_B08 | NQDIB154 | ACAGTTAAGC | PlateB_B08 | SQDIB154 | GAGTTGCAGG
PlateB_C08 | NQDIB155 | AGCCACACTA | PlateB_C08 | SQDIB155 | TCCACGAACA
PlateB_D08 | NQDIB156 | CAATCGTCTT | PlateB_D08 | SQDIB156 | TGACGGAGGA
PlateB_E08 | NQDIB157 | AGGAGCTTGT | PlateB_E08 | SQDIB157 | AATGAGTACG
PlateB_F08 | NQDIB158 | TTGAGCGGAG | PlateB_F08 | SQDIB158 | CGTCTTCCGA
PlateB_G08 | NQDIB159 | AGTAGCTCTC | PlateB_G08 | SQDIB159 | GACAGAGATT
PlateB_H08 | NQDIB160 | CACGCTGTCA | PlateB_H08 | SQDIB160 | TTACGCTAAC
PlateB_A09 | NQDIB161 | AAGACCTCTT | PlateB_A09 | SQDIB161 | CTCCTCGAAG
PlateB_B09 | NQDIB162 | GACCTCTTCT | PlateB_B09 | SQDIB162 | ATACCGCAGA
PlateB_C09 | NQDIB163 | TACTTCCTTG | PlateB_C09 | SQDIB163 | CCTATCTGAT
PlateB_D09 | NQDIB164 | TGCGATACGC | PlateB_D09 | SQDIB164 | GATCGGTTAC
PlateB_E09 | NQDIB165 | GCAGGCTTAA | PlateB_E09 | SQDIB165 | TGGTGAGGTG
PlateB_F09 | NQDIB166 | TAAGCTTGTG | PlateB_F09 | SQDIB166 | AACCGGCGTA
PlateB_G09 | NQDIB167 | ATGGTCCGCT | PlateB_G09 | SQDIB167 | AATACCGATC
PlateB_H09 | NQDIB168 | ATGTCAGAAG | PlateB_H09 | SQDIB168 | CGATACTCAA
PlateB_A10 | NQDIB169 | GACGAAGGTC | PlateB_A10 | SQDIB169 | GTAAGGCGGT
PlateB_B10 | NQDIB170 | ATCACCGTGA | PlateB_B10 | SQDIB170 | TTCAAGGTCG
PlateB_C10 | NQDIB171 | GCTACAGTGT | PlateB_C10 | SQDIB171 | TATCCGAGTA
PlateB_D10 | NQDIB172 | CGTCGAATAT | PlateB_D10 | SQDIB172 | AGCGCGCTTA
PlateB_E10 | NQDIB173 | CAACCATCGG | PlateB_E10 | SQDIB173 | CCGGAGACAT
PlateB_F10 | NQDIB174 | CGGTCCATTC | PlateB_F10 | SQDIB174 | GAGATAACTG
PlateB_G10 | NQDIB175 | AGAAGAGCCA | PlateB_G10 | SQDIB175 | TTGTAAGCGC
PlateB_H10 | NQDIB176 | CTATGCAATG | PlateB_H10 | SQDIB176 | CAAGAGGAGG
PlateB_A11 | NQDIB177 | CACTGAACCG | PlateB_A11 | SQDIB177 | AACCTTAGGA
PlateB_B11 | NQDIB178 | TACTGTGTGA | PlateB_B11 | SQDIB178 | CTGGCAACTC
PlateB_C11 | NQDIB179 | GCATTCTGTT | PlateB_C11 | SQDIB179 | GAACTTGTTG
PlateB_D11 | NQDIB180 | CTCCGCTAAG | PlateB_D11 | SQDIB180 | TGTGCAAGAT
PlateB_E11 | NQDIB181 | TCGCTTGAGA | PlateB_E11 | SQDIB181 | AATCGAGAGA
PlateB_F11 | NQDIB182 | AACTAGCCTT | PlateB_F11 | SQDIB182 | AGCGTGTCAG
PlateB_G11 | NQDIB183 | TTCGCTCAGG | PlateB_G11 | SQDIB183 | CTTGGTGATT
PlateB_H11 | NQDIB184 | CTCTACAACA | PlateB_H11 | SQDIB184 | GAAGCAGCAA
PlateB_A12 | NQDIB185 | TGAGTGTGTT | PlateB_A12 | SQDIB185 | TTCCGTCGAC
PlateB_B12 | NQDIB186 | TAGTTAGTCG | PlateB_B12 | SQDIB186 | CGAGATGCCA
PlateB_C12 | NQDIB187 | GCCTGATCCT | PlateB_C12 | SQDIB187 | AAGTTCGTGC
PlateB_D12 | NQDIB188 | CGAGTACAGG | PlateB_D12 | SQDIB188 | CGTCCATAAG
PlateB_E12 | NQDIB189 | GCCTAGATTA | PlateB_E12 | SQDIB189 | TTGTGGCATA
PlateB_F12 | NQDIB190 | TCGGCACTGT | PlateB_F12 | SQDIB190 | AGATCGGAAT
PlateB_G12 | NQDIB191 | CCGTGCAAGA | PlateB_G12 | SQDIB191 | CATTCTACTG
PlateB_H12 | NQDIB192 | CTGGCTGGTT | PlateB_H12 | SQDIB192 | ATCGCCGTAG

## Index Plates
Layout of the pierceable QIAseq 96-Unique Dual Index Set Adapter Plates for adapter ligation
QIAseq 96-Unique Dual Index Set A (96) |  | QIAseq 96-Unique Dual Index Set B (96)
Cat. no. 333725 |  | Cat. no. 333735
| QIAseq 96-Unique Dual Index Set A (96) layout (NQDIB001-NQDIB096). |  | QIAseq 96-Unique Dual Index Set B (96) layout (NQDIB097-NQDIB192).
Layout of the QIAseq 96-Unique Dual Index Set Index Plates in aluminum pouch for universal PCR
Each well contains one pre‑dispensed, dried sample index primer, and universal primer pair for a single reaction. In the universal PCR step, UDIS-96#K Index Plate Set A
or B must be paired with the appropriate UDIN-96#A Adapter Plate Set A or B used in the adapter ligation step, respectively.
QIAseq 96-Unique Dual Index Set A (96) |  | QIAseq 96-Unique Dual Index Set B (96)
Cat. no. 333725 |  | Cat. no. 333735
| QIAseq 96-Unique Dual Index Set A (96) layout (SQDIB001- SQDIB096). |  | QIAseq 96-Unique Dual Index Set B (96) layout (SQDIB097- SQDIB192).

## Dilute Lib
How to Dilute QIAseq Targeted DNA Panel Libraries
Experiment: |  |  | Initials: |  | Date:
Diluting the QIAseq DNA Panel Libraries to be loaded onto an Illumina instrument:
| Sample | pM | Dilution Factor | Lib Final conc (pM) | Volume H20 or HT1 Buffer (μl)
4000
| 1 |  | 0 | 3 | -3
| 2 |  | 0 | 3 | -3
| 3 |  | 0 | 3 | -3
| 4 |  | 0 | 3 | -3
| 5 |  | 0 | 3 | -3
| 6 |  | 0 | 3 | -3
| 7 |  | 0 | 3 | -3
| 8 |  | 0 | 3 | -3
| 9 |  | 0 | 3 | -3
| 10 |  | 0 | 3 | -3
| 11 |  | 0 | 3 | -3
| 12 |  | 0 | 3 | -3
| 13 |  | 0 | 3 | -3
| 14 |  | 0 | 3 | -3
| 15 |  | 0 | 3 | -3
| 16 |  | 0 | 3 | -3
| 17 |  | 0 | 3 | -3
| 18 |  | 0 | 3 | -3
| 19 |  | 0 | 3 | -3
| 20 |  | 0 | 3 | -3
| 21 |  | 0 | 3 | -3
| 22 |  | 0 | 3 | -3
| 23 |  | 0 | 3 | -3
| 24 |  | 0 | 3 | -3
| 25 |  | 0 | 3 | -3
| 26 |  | 0 | 3 | -3
| 27 |  | 0 | 3 | -3
| 28 |  | 0 | 3 | -3
| 29 |  | 0 | 3 | -3
| 30 |  | 0 | 3 | -3
| 31 |  | 0 | 3 | -3
| 32 |  | 0 | 3 | -3
| 33 |  | 0 | 3 | -3
| 34 |  | 0 | 3 | -3
| 35 |  | 0 | 3 | -3
| 36 |  | 0 | 3 | -3
| 37 |  | 0 | 3 | -3
| 38 |  | 0 | 3 | -3
| 39 |  | 0 | 3 | -3
| 40 |  | 0 | 3 | -3
| 41 |  | 0 | 3 | -3
| 42 |  | 0 | 3 | -3
| 43 |  | 0 | 3 | -3
| 44 |  | 0 | 3 | -3
| 45 |  | 0 | 3 | -3
| 46 |  | 0 | 3 | -3
| 47 |  | 0 | 3 | -3
| 48 |  | 0 | 3 | -3
| 49 |  | 0 | 3 | -3
| 50 |  | 0 | 3 | -3
| 51 |  | 0 | 3 | -3
| 52 |  | 0 | 3 | -3
| 53 |  | 0 | 3 | -3
| 54 |  | 0 | 3 | -3
| 55 |  | 0 | 3 | -3
| 56 |  | 0 | 3 | -3
| 57 |  | 0 | 3 | -3
| 58 |  | 0 | 3 | -3
| 59 |  | 0 | 3 | -3
| 60 |  | 0 | 3 | -3
| 61 |  | 0 | 3 | -3
| 62 |  | 0 | 3 | -3
| 63 |  | 0 | 3 | -3
| 64 |  | 0 | 3 | -3
| 65 |  | 0 | 3 | -3
| 66 |  | 0 | 3 | -3
| 67 |  | 0 | 3 | -3
| 68 |  | 0 | 3 | -3
| 69 |  | 0 | 3 | -3
| 70 |  | 0 | 3 | -3
| 71 |  | 0 | 3 | -3
| 72 |  | 0 | 3 | -3
| 73 |  | 0 | 3 | -3
| 74 |  | 0 | 3 | -3
| 75 |  | 0 | 3 | -3
| 76 |  | 0 | 3 | -3
| 77 |  | 0 | 3 | -3
| 78 |  | 0 | 3 | -3
| 79 |  | 0 | 3 | -3
| 80 |  | 0 | 3 | -3
| 81 |  | 0 | 3 | -3
| 82 |  | 0 | 3 | -3
| 83 |  | 0 | 3 | -3
| 84 |  | 0 | 3 | -3
| 85 |  | 0 | 3 | -3
| 86 |  | 0 | 3 | -3
| 87 |  | 0 | 3 | -3
| 88 |  | 0 | 3 | -3
| 89 |  | 0 | 3 | -3
| 90 |  | 0 | 3 | -3
| 91 |  | 0 | 3 | -3
| 92 |  | 0 | 3 | -3
| 93 |  | 0 | 3 | -3
| 94 |  | 0 | 3 | -3
| 95 |  | 0 | 3 | -3
| 96 |  | 0 | 3 | -3
Procedure for diluting and preparing the pooled library
[_] | 1. Dilute the sample libraries down to 4 nM using the calculations above. Add 3 μl of the library
sample to the appropriate volume of water or HT1 buffer.
[_] | 2. Multiple samples are now combined to create the "Lane Pool" – combined lane at 4 nM.
Therefore, take 5 μl of each sample and combine them in a tube. Mix thoroughly.
This is a safe stopping point.

## Denat and Dilut Lib MiSeq
How to Denature and Dilute Libraries for the MiSeq
Experiment: |  |  | Initials: |  | Date:
MiSeq Run Information: |  |  |  |  | Date:
Denaturing and diluting the sample pool libraries to be loaded onto the
MiSeq instrument
Preparation of fresh 0.2 N NaOH
1. Prepare fresh 0.2 N NaOH from 10 N NaOH stock as follows:
[_] | a. Add 20 μl of 10 N NaOH stock to 180 μl of nuclease-free water. Mix thoroughly.
The NaOH is now at a concentration of 1 N NaOH.
[_] | b. Add 100 μl of 1 N NaOH dilution from the step above to 400 μl of nuclease-free water.
Mix thoroughly. The NaOH is now at a concentration of 0.2 N NaOH.
Procedure for denaturing and diluting the pooled library into position #17
If loading one QIAseq Targeted DNA Panel pool
[_] | 1. Take 5 μl of the combined pooled 4 nM lane from the Dilute Lib tab and add it to
5 μl of 0.2 N NaOH. The samples are now at a concentration of 2 nM.
[_] | 2. Vortex briefly to mix, and then centrifuge the sample briefly.
[_] | 3. Denature the 2 nM sample for 5 min at room temperature.
[_] | 4. Add 990 μl of the HT1 Hybridization Buffer to the 10 μl reaction from step 3. The samples
are now at a 20 pM concentration and are ready to be diluted prior to being loaded onto the
MiSeq Reagent Cartridge.
[_] | 5. Dilute the sample again to a final concentration appropriate for loading onto the MiSeq
Reagent Cartridge. We typically load a final concentration of 10 pM onto the MiSeq
Reagent Cartridge in sample position #17. Hence, add 300 µl of HT1 Buffer to 300 µl of the
20 pM denatured library. Mix thoroughly before adding to position #17.
[_] | 6. Discard the unused 0.2 N NaOH.
If loading two different-sized QIAseq Targeted DNA Panel pools
[_] | 1. If combining libraries with the same number of primers, pool equal volumes of the individual
libraries at 4 nM together. If the libraries have different primer numbers, then combine the
libraries at volume ratios according to their number of primers. For example, Library A has
5000 primers at 4 nM and Library B has 600 primers at 4 nM, combining 50 µl Library A
with 6 µl Library B will result in similar coverage depth for both libraries A and B in the same
sequencing run.
Note: Fill in the number of samples for each of the DNA panels in cells B58–B61
that are going to be sequenced on the same sequencing run. Then pool the volume of
each 4 nM library (µl) as stated below.
In the drop-down menu in cell E57, select what instrument and reagents are going to be used
for the sequencing run.
Panel | Number of samples | Reads (M) | Percentage (%) | Data output (M) |  | Volume of each 4 nM library (µl)
Instrument_Reagent
1 |  | 0 | #DIV/0! | #DIV/0! |  | 0
2 |  | 0 | #DIV/0! | #DIV/0! |  | 0
3 |  | 0 | #DIV/0! | #DIV/0! |  | 0
4 |  | 0 | #DIV/0! | #DIV/0! |  | 0
| Sum | 0 |  | #N/A
[_] | 2. Combine the DNA pools accordingly, and mix thoroughly.
[_] | 3. Take 5 μl of the combined pooled 4 nM lane from the Dilute Lib tab and add it to
5 μl of 0.2 N NaOH. The samples are now at a concentration of 2 nM.
[_] | 4. Vortex briefly to mix, and then centrifuge the sample briefly.
[_] | 5. Denature the 2 nM sample for 5 min at room temperature.
[_] | 6. Add 990 μl of the HT1 Hybridization Buffer to the 10 μl reaction from step 5. The samples
are now at a 20 pM concentration and are ready to be diluted prior to being loaded onto the
MiSeq Reagent Cartridge.
[_] | 7. Dilute the sample again to a final concentration appropriate to load onto the MiSeq Reagent
Cartridge. We typically load a final concentration of 10 pM onto the MiSeq Reagent
Cartridge in sample position #17. Hence, add 300 µl of HT1 Buffer to 300 µl of the 20 pM
denatured library. Mix thoroughly before adding to position #17.
[_] | 8. Discard the unused 0.2 N NaOH.
Procedure for diluting and loading the Custom Read 1 Primer into position #18
Because we are using a Custom Read 1 Primer for the MiSeq run, we have to do the following:
Prepare the Custom Read 1 Primer – QIAseq A Read 1 Primer I as follows:
[_] | 1. Use HT1 Buffer to dilute the custom primer to yield 600 µl at a final concentration of 0.5 µM
by adding 3 µl of the QIAseq A Read 1 Primer I Custom Primer to 597 µl of HT1 Buffer.
How to add the Custom Read 1 Primer to the MiSeq Reagent Cartridge:
[_] | 1. Add 600 µl of the diluted QIAseq A Read 1 Primer I Custom Primer from above to position
#18 in the MiSeq Reagent Cartridge.
Important: Select Custom Primer for READ 1 in the MiSeq Set up. Also, no PhiX has to be added, because it will not sequence with the custom primer.
MiSeq Run Setup Information:
| 1. Sequencing Setup: |  | Select category Other > FASTQ Only
Read Type: Paired End Read
Cycles: 149 bp for R1, 149 bp for R2
CUSTOM READ 1 Primer
Indices: 10 bp

## Dilut and Denat Lib MiniSeq (hidden)
How to Denature and Dilute Libraries for the MiniSeq
Experiment: |  |  | Initials: |  | Date:
MiniSeq Run Information: |  |  |  |  | Date:
Denaturing and diluting the sample pool libraries to be loaded onto the
MiniSeq instrument
Preparation of fresh 0.1 N NaOH
1. Prepare Fresh 0.1 N NaOH from 10 N NaOH stock as follows:.
[_] | a. Add 10 μl of 10 N NaOH Stock to 90 μl of Nuclease-free Water. Mix thoroughly.
The NaOH is now at a concentration of 1 N NaOH.
[_] | b. Add 50 μl of 1 N NaOH Dilution from step a above to 450 μl of Nuclease-free Water.
Mix thoroughly. The NaOH is now at a concentration of 0.1 N NaOH.
Preparation of 200 mM Tris-HCl
[_] | 1. Dilute 100 µl of the 1 M Tris-HCl, pH 7.0 Stock to 400 µl if water.
[_] | 2. Mix thoroughly. The Tris-HCl is now at a concentration of 200 mM Tris-HCl.
Procedure for denaturing and diluting the pooled library into position #16
If loading one QIAseq Targeted DNA Panel pool
[_] | 1. Take 5 μl of the combined pooled 4 nM lane from the Dilute Lib tab and add it to
5 μl of 0.1 N NaOH. The samples are now at a concentration of 2 nM.
[_] | 2. Vortex briefly to mix, and then centrifuge the sample briefly.
[_] | 3. Denature the 2 nM sample for 5 min at room temperature.
[_] | 4. Add 5 μl of the 200 mM Tris-HCl, pH 7 to the 10 μl reaction from step 3.
[_] | 5. Vortex briefly to mix and then centrifuge the sample briefly.
[_] | 6. Add 985 μl of the HT1 Hybridization Buffer to the 15 μl reaction from step 5. The Samples are
now at a 20 pM concentration and are ready to be diluted prior to being loaded onto the MiniSeq
Reagent Cartridge.
[_] | 7. Dilute the sample again to a final concentration appropriate to load onto the MiniSeq Reagent
Cartridge. We typically load a final concentration of 1.2 pM onto the MiniSeq Reagent Cartridge
in Sample Position #16. Hence add 470 µl of HT1 Buffer to 30 µl of the 20 pM denatured library
for a final volume of 500 µl. Mix thoroughly before adding to position #16.
[_] | 8. Discard the unused 0.1 N NaOH.
If loading two Different sized QIAseq Targeted DNA Panel pools
[_] | 1. If combining libraries with the same number of primers, pool equal volumes of the individual
libraries at 4 nM together. If the libraries have different primer numbers, then combine the
libraries at volume ratios according to their number of primers. For example, Library A has
5000 primers at 4 nM, and Library B has 600 primers at 4 nM; combining 50 µl Library A with
6 µl Library B will result in similar coverage depth for both libraries A and B in the same
sequencing run.
Note: Fill in the Number of Samples for each of the DNA Panels in Cells B58 - B61,
that are going to be sequenced on the same sequencing run. Then pool the volume of
each 4 nM library (µl) as stated below.
In the drop-down menu in Cell E57 select what Instrument and Reagents are going to be used
for the sequencing run.
Panel | Number of Samples | Reads (M) | Percentage (%) | Data Output (M) |  | Volume of each 4 nM Library (µl)
Instrument_Reagent
1 |  | 0 | #DIV/0! | #DIV/0! |  | 0
2 |  | 0 | #DIV/0! | #DIV/0! |  | 0
3 |  | 0 | #DIV/0! | #DIV/0! |  | 0
4 |  | 0 | #DIV/0! | #DIV/0! |  | 0
| Sum | 0 |  | #N/A
[_] | 2. Combine the DNA pools accordingly, and mix thoroughly.
[_] | 3. Take 5 μl of the combined pooled 4 nM lane from the Dilute Lib tab and add it to
5 μl of 0.1 N NaOH. The samples are now at a concentration of 2 nM.
[_] | 4. Vortex briefly to mix, and then centrifuge the sample briefly.
[_] | 5. Denature the 2 nM sample for 5 min at room temperature.
[_] | 6. Add 5 μl of the 200 mM Tris-HCl, pH 7 to the 10 μl reaction from step 5.
[_] | 7. Vortex briefly to mix and then centrifuge the sample briefly.
[_] | 8. Add 985 μl of the HT1 Hybridization Buffer to the 15 μl reaction from step 7. The Samples are
now at a 20 pM concentration and are ready to be diluted prior to being loaded onto the MiniSeq
Reagent Cartridge.
[_] | 9. Dilute the sample again to a final concentration appropriate to load onto the MiniSeq Reagent
Cartridge. We typically load a final concentration of 1.2 pM onto the MiniSeq Reagent Cartridge
in Sample Position #16. Hence add 470 µl of HT1 Buffer to 30 µl of the 20 pM denatured library
for a final volume of 500 µl. Mix thoroughly before adding to position #16.
#10.
[_] | 10. Discard the unused 0.1 N NaOH.
Procedure for diluting and loading the Custom Read 1 Primer into position #15
Because we are using a Custom Read 1 Primer for the MiniSeq run, we have to do the following:
Prepare the Custom Read 1 Primer – QIAseq A Read 1 Primer I as follows:
[_] | 1. Use HT1 Buffer to dilute the custom primer to yield 1 ml at a final concentration of 0.3 µM by
adding 3 µl of the QIAseq A Read 1 Primer I Custom Primer to 997 µl of HT1 Buffer.
How to add the Custom Read 1 Primer to the MiniSeq Reagent Cartridge:
[_] | 1. Add 1 ml of the diluted QIAseq A Read 1 Primer I Custom Primer from above to position #15
in the MiniSeq Reagent Cartridge.
Important: Select Custom Primer for READ 1 in the MiniSeq Set up. Also, no PhiX has to be added, because it will not sequence with the custom primer.
MiniSeq Run Setup Information:
| 1. Sequencing Setup: |  | Select category Other > FASTQ Only
Read Type: Paired End Read
Cycles: 149 bp for R1, 149 bp for R2
CUSTOM READ 1 Primer
Indices: 10 bp

## Dilut and Denat Lib NextSeq (hidden)
How to Denature and Dilute Libraries for the NextSeq
Experiment: |  |  | Initials: |  | Date:
NextSeq Run Information: |  |  |  |  | Date:
Denaturing and diluting the sample pool libraries to be loaded onto the
NextSeq instrument
Preparation of fresh 0.2 N NaOH
1. Prepare Fresh 0.2 N NaOH from 10 N NaOH stock as follows:.
[_] | a. Add 20 μl of 10 N NaOH Stock to 180 μl of Nuclease-free Water. Mix thoroughly.
The NaOH is now at a concentration of 1 N NaOH.
[_] | b. Add 100 μl of 1 N NaOH Dilution from step a above to 400 μl of Nuclease-free Water.
Mix thoroughly. The NaOH is now at a concentration of 0.2 N NaOH.
Preparation of 200 mM Tris-HCl
[_] | 1. Dilute 100 µl of the 1 M Tris-HCl, pH 7.0 Stock to 400 µl if water.
[_] | 2. Mix thoroughly. The Tris-HCl is now at a concentration of 200 mM Tris-HCl.
Procedure for denaturing and diluting the pooled library into position #10
If loading one QIAseq Targeted DNA Panel pool
[_] | 1. Take 5 μl of the combined pooled 4 nM lane from the Dilute Lib tab and add it to
5 μl of 0.2 N NaOH. The samples are now at a concentration of 2 nM.
[_] | 2. Vortex briefly to mix, and then centrifuge the sample briefly.
[_] | 3. Denature the 2 nM sample for 5 min at room temperature.
[_] | 4. Add 5 μl of the 200 mM Tris-HCl, pH 7 to the 10 μl reaction from step 3.
[_] | 5. Vortex briefly to mix and then centrifuge the sample briefly.
[_] | 6. Add 985 μl of the HT1 Hybridization Buffer to the 15 μl reaction from step 5. The Samples are
now at a 20 pM concentration and are ready to be diluted prior to being loaded onto the
NextSeq Reagent Cartridge.
[_] | 7. Dilute the sample again to a final concentration appropriate to load onto the NextSeq Reagent
Cartridge. We typically load a final concentration of 1.2 pM onto the NextSeq Reagent Cartridge
in Sample Position #10. Hence add two aliquots 611 µl (1.22 ml) of HT1 Buffer to 78 µl of the
20 pM denatured library for a final volume of 1.3 ml. Mix thoroughly before adding to position
#10.
[_] | 8. Discard the unused 0.2 N NaOH.
If loading two Different sized QIAseq Targeted DNA Panel pools
[_] | 1. If combining libraries with the same number of primers, pool equal volumes of the individual
libraries at 4 nM together. If the libraries have different primer numbers, then combine the
libraries at volume ratios according to their number of primers. For example, Library A has
5000 primers at 4 nM, and Library B has 600 primers at 4 nM; combining 50 µl Library A with
6 µl Library B will result in similar coverage depth for both libraries A and B in the same
sequencing run.
Note: Fill in the Number of Samples for each of the DNA Panels in Cells B58 - B61,
that are going to be sequenced on the same sequencing run. Then pool the volume of
each 4 nM library (µl) as stated below.
In the drop-down menu in Cell E57 select what Instrument and Reagents are going to be used
for the sequencing run.
Panel | Number of Samples | Reads (M) | Percentage (%) | Data Output (M) |  | Volume of each 4 nM Library (µl)
Instrument_Reagent
1 |  | 0 | #DIV/0! | #DIV/0! |  | 0
2 |  | 0 | #DIV/0! | #DIV/0! |  | 0
3 |  | 0 | #DIV/0! | #DIV/0! |  | 0
4 |  | 0 | #DIV/0! | #DIV/0! |  | 0
| Sum | 0 |  | #N/A
[_] | 2. Combine the DNA pools accordingly, and mix thoroughly.
[_] | 3. Take 5 μl of the combined pooled 4 nM lane from the Dilute Lib tab and add it to
5 μl of 0.2 N NaOH. The samples are now at a concentration of 2 nM.
[_] | 4. Vortex briefly to mix, and then centrifuge the sample briefly.
[_] | 5. Denature the 2 nM sample for 5 min at room temperature.
[_] | 6. Add 5 μl of the 200 mM Tris-HCl, pH 7 to the 10 μl reaction from step 5.
[_] | 7. Vortex briefly to mix and then centrifuge the sample briefly.
[_] | 8. Add 985 μl of the HT1 Hybridization Buffer to the 15 μl reaction from step 7. The Samples are
now at a 20 pM concentration and are ready to be diluted prior to being loaded onto the
NextSeq Reagent Cartridge.
[_] | 9. Dilute the sample again to a final concentration appropriate to load onto the NextSeq Reagent
Cartridge. We typically load a final concentration of 1.2 pM onto the NextSeq Reagent Cartridge
in Sample Position #10. Hence add two aliquots 611 µl (1.22 ml) of HT1 Buffer to 78 µl of the
20 pM denatured library for a final volume of 1.3 ml. Mix thoroughly before adding to position
#10.
[_] | 10. Discard the unused 0.2 N NaOH.
Procedure for diluting and loading the Custom Read 1 Primer into position #7
Because we are using a Custom Read 1 Primer for the NextSeq run, we have to do the following:
Prepare the Custom Read 1 Primer – QIAseq A Read 1 Primer I as follows:
[_] | 1. Use HT1 Buffer to dilute the custom primer to yield 2 ml at a final concentration of 0.3 µM by
adding 6 µl of the QIAseq A Read 1 Primer I Custom Primer to two aliquots of 997 µl of
HT1 Buffer.
How to add the Custom Read 1 Primer to the NextSeq Reagent Cartridge:
[_] | 1. Add 2 ml of the diluted QIAseq A Read 1 Primer I Custom Primer from above to position #7
in the NextSeq Reagent Cartridge.
Important: Select Custom Primer for READ 1 in the NextSeq Set up. Also, no PhiX has to be added, because it will not sequence with the custom primer.
NextSeq Run Setup Information:
| 1. Sequencing Setup: |  | Select category Other > FASTQ Only
Read Type: Paired End Read
Cycles: 149 bp for R1, 149 bp for R2
CUSTOM READ 1 Primer
Indices: 10 bp

## LRMv3 Import Sample Sheet
[Header]
Experiment Name
Date
Module | GenerateFASTQ - 3.0.1
Workflow | GenerateFASTQ
Library Prep Kit | QIAseq Targeted DNA Panels UDIs
Index Kit | QIAseq UDIs
Description
Chemistry | Amplicon
[Reads]
149
149
[Settings]
adapter | CTGTCTCTTATACACATCT
customread1primermix | C1
[Data]
Sample_ID | Description | Index_Plate_Well | I7_Index_ID | index | I5_Index_ID | index2 | Sample_Project
1 | _ |  | #N/A | #N/A | #N/A | #N/A
2 | _ |  | #N/A | #N/A | #N/A | #N/A
3 | _ |  | #N/A | #N/A | #N/A | #N/A
4 | _ |  | #N/A | #N/A | #N/A | #N/A
5 | _ |  | #N/A | #N/A | #N/A | #N/A
6 | _ |  | #N/A | #N/A | #N/A | #N/A
7 | _ |  | #N/A | #N/A | #N/A | #N/A
8 | _ |  | #N/A | #N/A | #N/A | #N/A
9 | _ |  | #N/A | #N/A | #N/A | #N/A
10 | _ |  | #N/A | #N/A | #N/A | #N/A
11 | _ |  | #N/A | #N/A | #N/A | #N/A
12 | _ |  | #N/A | #N/A | #N/A | #N/A
13 | _ |  | #N/A | #N/A | #N/A | #N/A
14 | _ |  | #N/A | #N/A | #N/A | #N/A
15 | _ |  | #N/A | #N/A | #N/A | #N/A
16 | _ |  | #N/A | #N/A | #N/A | #N/A
17 | _ |  | #N/A | #N/A | #N/A | #N/A
18 | _ |  | #N/A | #N/A | #N/A | #N/A
19 | _ |  | #N/A | #N/A | #N/A | #N/A
20 | _ |  | #N/A | #N/A | #N/A | #N/A
21 | _ |  | #N/A | #N/A | #N/A | #N/A
22 | _ |  | #N/A | #N/A | #N/A | #N/A
23 | _ |  | #N/A | #N/A | #N/A | #N/A
24 | _ |  | #N/A | #N/A | #N/A | #N/A
25 | _ |  | #N/A | #N/A | #N/A | #N/A
26 | _ |  | #N/A | #N/A | #N/A | #N/A
27 | _ |  | #N/A | #N/A | #N/A | #N/A
28 | _ |  | #N/A | #N/A | #N/A | #N/A
29 | _ |  | #N/A | #N/A | #N/A | #N/A
30 | _ |  | #N/A | #N/A | #N/A | #N/A
31 | _ |  | #N/A | #N/A | #N/A | #N/A
32 | _ |  | #N/A | #N/A | #N/A | #N/A
33 | _ |  | #N/A | #N/A | #N/A | #N/A
34 | _ |  | #N/A | #N/A | #N/A | #N/A
35 | _ |  | #N/A | #N/A | #N/A | #N/A
36 | _ |  | #N/A | #N/A | #N/A | #N/A
37 | _ |  | #N/A | #N/A | #N/A | #N/A
38 | _ |  | #N/A | #N/A | #N/A | #N/A
39 | _ |  | #N/A | #N/A | #N/A | #N/A
40 | _ |  | #N/A | #N/A | #N/A | #N/A
41 | _ |  | #N/A | #N/A | #N/A | #N/A
42 | _ |  | #N/A | #N/A | #N/A | #N/A
43 | _ |  | #N/A | #N/A | #N/A | #N/A
44 | _ |  | #N/A | #N/A | #N/A | #N/A
45 | _ |  | #N/A | #N/A | #N/A | #N/A
46 | _ |  | #N/A | #N/A | #N/A | #N/A
47 | _ |  | #N/A | #N/A | #N/A | #N/A
48 | _ |  | #N/A | #N/A | #N/A | #N/A
49 | _ |  | #N/A | #N/A | #N/A | #N/A
50 | _ |  | #N/A | #N/A | #N/A | #N/A
51 | _ |  | #N/A | #N/A | #N/A | #N/A
52 | _ |  | #N/A | #N/A | #N/A | #N/A
53 | _ |  | #N/A | #N/A | #N/A | #N/A
54 | _ |  | #N/A | #N/A | #N/A | #N/A
55 | _ |  | #N/A | #N/A | #N/A | #N/A
56 | _ |  | #N/A | #N/A | #N/A | #N/A
57 | _ |  | #N/A | #N/A | #N/A | #N/A
58 | _ |  | #N/A | #N/A | #N/A | #N/A
59 | _ |  | #N/A | #N/A | #N/A | #N/A
60 | _ |  | #N/A | #N/A | #N/A | #N/A
61 | _ |  | #N/A | #N/A | #N/A | #N/A
62 | _ |  | #N/A | #N/A | #N/A | #N/A
63 | _ |  | #N/A | #N/A | #N/A | #N/A
64 | _ |  | #N/A | #N/A | #N/A | #N/A
65 | _ |  | #N/A | #N/A | #N/A | #N/A
66 | _ |  | #N/A | #N/A | #N/A | #N/A
67 | _ |  | #N/A | #N/A | #N/A | #N/A
68 | _ |  | #N/A | #N/A | #N/A | #N/A
69 | _ |  | #N/A | #N/A | #N/A | #N/A
70 | _ |  | #N/A | #N/A | #N/A | #N/A
71 | _ |  | #N/A | #N/A | #N/A | #N/A
72 | _ |  | #N/A | #N/A | #N/A | #N/A
73 | _ |  | #N/A | #N/A | #N/A | #N/A
74 | _ |  | #N/A | #N/A | #N/A | #N/A
75 | _ |  | #N/A | #N/A | #N/A | #N/A
76 | _ |  | #N/A | #N/A | #N/A | #N/A
77 | _ |  | #N/A | #N/A | #N/A | #N/A
78 | _ |  | #N/A | #N/A | #N/A | #N/A
79 | _ |  | #N/A | #N/A | #N/A | #N/A
80 | _ |  | #N/A | #N/A | #N/A | #N/A
81 | _ |  | #N/A | #N/A | #N/A | #N/A
82 | _ |  | #N/A | #N/A | #N/A | #N/A
83 | _ |  | #N/A | #N/A | #N/A | #N/A
84 | _ |  | #N/A | #N/A | #N/A | #N/A
85 | _ |  | #N/A | #N/A | #N/A | #N/A
86 | _ |  | #N/A | #N/A | #N/A | #N/A
87 | _ |  | #N/A | #N/A | #N/A | #N/A
88 | _ |  | #N/A | #N/A | #N/A | #N/A
89 | _ |  | #N/A | #N/A | #N/A | #N/A
90 | _ |  | #N/A | #N/A | #N/A | #N/A
91 | _ |  | #N/A | #N/A | #N/A | #N/A
92 | _ |  | #N/A | #N/A | #N/A | #N/A
93 | _ |  | #N/A | #N/A | #N/A | #N/A
94 | _ |  | #N/A | #N/A | #N/A | #N/A
95 | _ |  | #N/A | #N/A | #N/A | #N/A
96 | _ |  | #N/A | #N/A | #N/A | #N/A